TENDER ENQUIRY DOCUMENT

FOR PURCHASE OF MEDICAL EQUIPMENT

FOR

Tertiary Cancer Care Centre

at

Government Medical College, Kozhikode

HLL/ID/KKD-TCC/15-16/27

BY

HLL Lifecare Limited

(A GOVERNMENT OF INDIA ENTERPRISE)

Infrastructure Development Division

“Adarsh” TC 6/1781, Vettamukku,

Thirumala P.O

Thiruvananthapuram-695006

INDEX

Section	Topic                                                                                                                   Page No.

Section I	–	Notice inviting Tender (NIT) ------------------------------------------------------------	03

Section II	–	General Instructions to Tenderers (GIT) --------------------------------------------------	06

Section III 	–	Special Instructions to Tenderers (SIT) ----------------------------------------------------	25

Section IV	–	General Conditions of Contract (GCC) --------------------------------------------------	27

Section V	–	Special Conditions of Contract (SCC) ---------------------------------------------------	42

Section VI	–	List of Requirements -------------------------------------------------------------------	43

Section VII	–	Technical Specifications ---------------------------------------------------------------	45

Section VIII	–	Quality Control Requirements ---------------------------------------------------------	75

Section IX	–	Qualification Criteria -----------------------------------------------------------------	76

Section X	–	Tender Form --------------------------------------------------------------------------	79

Section XI	–	Price Schedules -----------------------------------------------------------------------	80

Section XII	–	Questionnaire ------------------------------------------------------------------------	84

Section XIII	–	Bank Guarantee Form for EMD --------------------------------------------------------	85

Section XIV	–	Manufacturer’s Authorisation Form -----------------------------------------------------	86

Section XV	–	Bank Guarantee Form for Performance Security /CMC Security --------------------------	87

Section XVI	–	Contract Form (A & B) ----------------------------------------------------------------	88

Section XVII	–	Proforma of Consignee Receipt Certificate ----------------------------------------------	92

Section XVIII–  Proforma of Final Acceptance Certificate by the Consignee -------------------------------	93

Section XIX	–	Instructions from Ministry of Shipping/Surface Transport (Annexure 1) -----------------	95

Section XX	–	Check List for the Tenderers -----------------------------------------------------------	     99

Section XXI	–	Consignee----------------------------------------------------------------------------	103

SECTION I

NOTICE INVITING TENDERS (NIT)

For Global Tender

HLL Lifecare Limited

(A GOVERNMENT OF INDIA ENTERPRISE)

- Infrastructure Development Division

- “Adarsh” TC 6/1781, Vettamukku,

- Thirumala P.O

- Thiruvananthapuram-695006

- FOR

Tertiary Cancer Care Centre

at

Government Medical College, Kozhikode

TenderEnquiryNo: HLL/ID/KKD-TCC/15-16/27                                                     Dated 24.08.2015

NOTICE INVITING TENDERS (NIT)

- Infrastructure Development Division of HLL Lifecare Ltd, for and on behalf of The principal, Government Medical College, Kozhikode invites sealed tenders, from eligible and qualified bidders for supply of following equipment for Tertiary Cancer Care Centre:

| Sch. No. | Equipment | Total Quantity | Amount of EMD (in Rs.) |
| --- | --- | --- | --- |
| 1 | Dual Head SPECT (CT) Gamma Camera including accessories. | 1 | 7,82,500 |

(2)	Tender No.: HLL/ID/KKD-TCC/15-16/27

| Sl No. | Description | Schedule |
| --- | --- | --- |
| i. | Dates of sale of tender enquiry documents | 25.08.2015 to 15.09.2015 |
| ii. | Place of sale of Tender Enquiry Documents | HLL Lifecare Ltd, (A Government of India Enterprise), Infrastructure Development Division, “Adarsh” TC 6/1781, Vettamukku, Thirumala P.O, Thiruvananthapuram-695006 |
| iii. | Cost of the Tender Enquiry Document | Rs. 5000/- |
| iv | Pre Tender Meeting Date & Time | 03.09.2015 at 11 a.m. |
| v | Pre bid conference Venue | Same as 2 (ii) |
| vi. | Closing date & time for receipt of Tender | 16.09.2015 at 15:00 hrs |
| vii. | Time and date of opening of Techno – Commercial tenders | 17.09.2015 at 15:00 hrs |
| viii | Venue of Opening of Techno Commercial Tender | Same as above |

3.	Interested bidders may obtain further information about this requirement from the above office selling the documents. Tender Enquiry Documents may be purchased on payment of non-refundable fee of Rs 5000/- per set in the form of account payee Demand Draft/Pay Order/Cashier’s Cheque/Banker’s Cheque, drawn on a scheduled bank in India, in favour of “HLL Lifecare Limited” payable at Thiruvananthapuram.

4.	Deleted

5.	Bidders may also download the tender enquiry documents from the web site, www. eprocure.gov.in/cppp or www.lifecarehll.com and submit its tender by utilizing the downloaded document, along with the required non-refundable fee as mentioned in Para 3 above. Amendment if any to Tender document shall be published in the website(s) only.

6.	All prospective tenderers may attend the Pre bid conference. The venue, date and time indicated in the Para 2 above.

7.	Bidders shall ensure that their tenders, complete in all respects, are submitted at Infrastructure Development Division, “Adarsh” TC 6/1781, Vettamukku, Thirumala P.O, Thiruvananthapuram-695006 on or before the closing date and time indicated in the Para 2 above, failing which the tenders will be treated as late and rejected. The tenders sent by post/courier must reach the above said address on or before the closing date & time indicated in Para2 above, failing which the tenders will be treated as late tender and rejected.

8.	In the event of any of the above mentioned dates being declared as a holiday / closed day for HLL, the tenders will be sold/received/opened on the next working day at the appointed time.

9.	The Tender Enquiry Documents are not transferable.

10. All tenders must be accompanied by EMD as mentioned against each item. The tenders without EMD shall be rejected.

11. The purchaser reserves the right to accept or reject a tender either in part or in full without assigning

any reason.

For and on behalf of The Principal,

Tertiary Cancer Care Centre

at-Government Medical College, Kozhikode

DVP (Tech)

HLL Lifecare Limited,

Infrastructure Development Division,

“Adarsh” TC 6/1781,

Vettamukku, Thirumala P.O,

Thiruvananthapuram-695006

SECTION - II

GENERAL INSTRUCTIONS TO TENDERERS (GIT)

CONTENTS

| Sl. No. | Topic | Page No. |
| --- | --- | --- |
| A | PREAMBLE |  |
| 1 | Definitions and Abbreviations | 8 |
| 2 | Introduction | 9 |
| 3 | Availability of Funds | 10 |
| 4 | Language of Tender | 10 |
| 5 | Eligible Tenderers | 10 |
| 6 | Eligible Goods and Services | 10 |
| 7 | Tendering Expense | 10 |
| B | TENDER ENQUIRY DOCUMENTS |  |
| 8 | Contents of Tender Enquiry Documents | 10 |
| 9 | Amendments to Tender Enquiry Documents | 11 |
| 10 | Clarification of Tender Enquiry Documents | 11 |
| C | PREPARATION OF TENDERS |  |
| 11 | Documents Comprising the Tender | 11 |
| 12 | Tender Currencies | 12 |
| 13 | Tender Prices | 13 |
| 14 | Indian Agent | 15 |
| 15 | Firm Price / Variable Price | 16 |
| 16 | Alternative Tenders | 16 |
| 17 | Documents Establishing Tenderer’s Eligibility and Qualifications | 16 |
| 18 | Documents Establishing Good’s Conformity to Tender Enquiry Document | 16 |
| 19 | Earnest Money Deposit (EMD) | 17 |
| 20 | Tender Validity | 17 |
| 21 | Signing and Sealing of Tender | 18 |

| D | SUBMISSION OF TENDERS |  |
| --- | --- | --- |
| 22 | Submission of Tenders | 18 |
| 23 | Late Tender | 19 |
| 24 | Alteration and Withdrawal of Tender | 19 |
| E | TENDER OPENING |  |
| 25 | Opening of Tenders | 19 |
| F | SCRUTINY AND EVALUATION OF TENDERS |  |
| 26 | Basic Principle | 20 |
| 27 | Preliminary Scrutiny of Tenders | 20 |
| 28 | Minor Infirmity/Irregularity/Non-Conformity | 20 |
| 29 | Discrepancy in Prices | 21 |
| 30 | Discrepancy between original and copies of Tender | 21 |
| 31 | Qualification Criteria | 21 |
| 32 | Conversion of Tender Currencies to Indian Rupees | 21 |
| 33 | Schedule-wise Evaluation | 21 |
| 34 | Comparison of Tenders | 21 |
| 35 | Additional Factors and Parameters for Evaluation and Ranking of Responsive Tenders | 22 |
| 36 | Tenderer’s capability to perform the contract | 22 |
| 37 | Contacting the Purchaser | 22 |
| G | AWARD OF CONTRACT |  |
| 38 | Purchaser’s Right to Accept any Tender and to Reject any or All Tenders | 22 |
| 39 | Award Criteria | 23 |
| 40 | Variation of Quantities at the Time of Award | 23 |
| 41 | Notification of Award | 23 |
| 42 | Issue of Contract | 23 |
| 43 | Non-receipt of Performance Security and Contract by the Purchaser/Consignee | 23 |
| 44 | Return of EMD | 23 |
| 45 | Publication of Tender Result | 24 |
| 46 | Corrupt or Fraudulent Practices | 24 |

SECTION – II

GENERAL INSTRUCTIONS TO TENDERERS (GIT)

A. PREAMBLE

1.	Definitions and Abbreviations

1.1	The following definitions and abbreviations, which have been used in these documents shall have the meanings as indicated below:

1.2.	Definitions:

(i)	“Purchaser” means the organization purchasing goods and services as incorporated in the Tender Enquiry document.

(ii)	“Tender” means Bids / Quotation / Tender received from a Firm / Tenderer / Bidder.

(iii)	“Tenderer” means Bidder/ the Individual or Firm submitting Bids / Quotation / Tender

(iii)	“Supplier” means the individual or the firm supplying the goods and services as incorporated in the contract.

(iv)	“Goods” means the articles, material, commodities, livestock, furniture, fixtures, raw material, spares, instruments, machinery, equipment, medical equipment, industrial plant etc. which the supplier is required to supply to the purchaser under the contract.

(v)	“Services” means services allied and incidental to the supply of goods, such as transportation, installation, commissioning, provision of technical assistance, training, after sales service, maintenance service and other such obligations of the supplier covered under the contract.

(vi)	“Earnest Money Deposit” (EMD) means Bid Security/ monetary or financial guarantee to be furnished by a tenderer along with its tender.

(vii)	“Contract” means the written agreement entered into between the purchaser and/or consignee and the supplier, together with all the

documents mentioned therein and including all attachments, annexure etc. therein.

(viii)	“Performance Security” means monetary or financial guarantee to be furnished by the successful tenderer for due performance of the contract placed on it. Performance Security is also known as Security Deposit.

(ix)	“Consignee” means the Hospital/Institute/Medical College/ person to whom the goods are required to be delivered as specified in the Contract. If the goods are required to be delivered to a person as an interim consignee for the purpose of despatch to another person as provided in the Contract then that “another” person is the consignee, also known as ultimate consignee.

(x)	“Specification” means the document/standard that prescribes the requirement with which goods or service has to conform.

(xi)	“Inspection” means activities such as measuring, examining, testing, gauging one or more characteristics of the product or service and comparing the same with the specified requirement to determine conformity.

(xii)	“Day” means calendar day.

1.3 	Abbreviations:

(i)	“T E Document” means Tender Enquiry Document

(ii)	“NIT” means Notice Inviting Tenders.

(iii)	“GIT” means General Instructions to Tenderers

(iv)	“SIT” means Special Instructions to Tenderers

(v)	“GCC” means General Conditions of Contract

(vi)	“SCC” means Special Conditions of Contract

(vii)	“DGS&D” means Directorate General of Supplies and Disposals

(viii)	“NSIC” means National Small Industries Corporation

(ix)	“PSU” means Public Sector Undertaking

(x)	“CPSU” means Central Public Sector Undertaking

(xi)	“LSI” means Large Scale Industry

(xii)	“SSI” means Small Scale Industry

(xiii)	“LC” means Letter of Credit

(xiv)	“DP” means Delivery Period

(xv)	“BG” means Bank Guarantee

(xvi)	“ED” means Excise Duty

(xvii)	“CD” means Custom Duty

(xviii)	“VAT” means Value Added Tax

(xix)	“CENVAT” means Central Value Added Tax

(xx)	“CST” means Central Sales Tax

(xxi)	“RR” means Railway Receipt

(xxii)	“BL” means Bill of Lading

(xxiii)	“FOB” means Free on Board

(xxiv)	“FCA” means Free Carrier

(xxv)	 “FOR” means Free On Rail

(xxvi)	 “CIF” means Cost, Insurance and Freight

(xxvii)	 “CIP (Destinations)” means Carriage and Insurance Paid up to named port of destination. Additionally the Insurance (local transportation and storage) would be extended and borne by the Supplier from ware house to the consignee site for a period including 3 months beyond date of delivery.

(xxviii)	“DDP” means Delivery Duty Paid named place of destination (consignee site)

(xxix)	“INCOTERMS” means International Commercial Terms as on the date of Tender Opening

(xxx)	”MOH&FW” means Ministry of Health & Family Welfare, Government of Kerala

- “CMC” means Comprehensive maintenance Contract (labour, spare and  preventive

maintenance)

- “RT” means Re-Tender.

- TCCC- Tertiary Cancer Care Centre, at Government Medical College, Kozhikode

2.	Introduction

2.1	The Purchaser has issued these TE documents for purchase of Medical equipment/ goods and related services as mentioned in Section – VI – “List of Requirements”, which also indicates, interalia, the required delivery schedule, terms and place of delivery.

2.2	This section (Section II - “General Instruction Tenderers”) provides the relevant information as well as instructions to assist the prospective tenderers in preparation and submission of tenders. It also includes the mode and procedure to be adopted by the purchaser for receipt and opening as well as scrutiny and evaluation of tenders and subsequent placement of contract.

2.3 	The tenderers shall also read the Special Instructions to Tenderers (SIT) related to this purchase, as contained in Section III of these documents and follow the same accordingly. Whenever there is a conflict between the GIT and the SIT, the provisions contained in the SIT shall prevail over those in the GIT.

2.4	Before formulating the tender and submitting the same to the purchaser, the tenderer should read and examine all the terms, conditions, instructions, checklist etc. contained in the TE documents. Failure to provide and/or comply with the required information, instructions etc. incorporated in these TE documents may result in rejection of its tender.

3.	Availability of Funds

3.1 	Deleted

4.	Language of Tender

4.1 	The tender submitted by the tenderer and all subsequent correspondence and documents relating to the tender exchanged between the tenderer and the purchaser, shall be written in the English language, unless otherwise specified in the Tender Enquiry. However, the language of any printed literature furnished by the tenderer in connection with its tender may be written in any other language provided the same is accompanied by an English translation and, for purposes of interpretation of the tender, the English translation shall prevail.

4.2 	The tender submitted by the tenderer and all subsequent correspondence and documents relating to the tender exchanged between the tenderer and the purchaser, may also be written in the Hindi language, provided that the same are accompanied by English translation, in which case, for purpose of interpretation of the tender etc, the English translations shall prevail.

5.	Eligible Tenderers

5.1 	This invitation for tenders is open to all suppliers who fulfil the eligibility criteria specified in these documents.

6.	Eligible Goods and Services

6.1	All goods and related services to be supplied under the contract shall have their origin in India or any other country with which India has not banned trade relations. The term “origin” used in this clause means the place where the goods are mined, grown, produced, or manufactured or from where the related services are arranged and supplied.

7.	Tendering Expense

7.1	The tenderer shall bear all costs and expenditure incurred and/or to be incurred by it in connection with its tender including preparation, mailing and submission of its tender and for subsequent processing the same. The purchaser will, in no case be responsible or liable for any such cost, expenditure etc regardless of the conduct or outcome of the tendering process.

B. TENDER ENQUIRY DOCUMENTS

This tender document should be read in conjunction with the Notice Inviting Tender (NIT) a copy of which is enclosed with this document. All clauses should be read in conjunction with any other instructions given else where in this document on the same subject matter of the clause.

8.	Content of Tender Enquiry Documents

8.1	In addition to Section I – “Notice inviting Tender” (NIT), the TE documents include:

- Section II	– General Instructions to Tenderers (GIT)

- Section III 	– Special Instructions to Tenderers (SIT)

- Section IV	– General Conditions of Contract (GCC)

- Section V	– Special Conditions of Contract (SCC)

- Section VI	– List of Requirements

- Section VII	– Technical Specifications

- Section VIII	– Quality Control Requirements

- Section IX	– Qualification Criteria

- Section X	– Tender Form

- Section XI	– Price Schedules

- Section XII	– Questionnaire

- Section XIII	– Bank Guarantee Form for EMD

- Section XIV	– Manufacturer’s Authorisation Form

- Section XV	– Bank Guarantee Form for Performance Security/CMC Security

- Section XVI	– Contract Forms A & B

- Section XVII	– Proforma of Consignee Receipt Certificate

- Section XVIII	– Proforma of Final Acceptance Certificate by the consignee

- Section XIX	–	Instructions from Ministry of Shipping/ Surface Transport (Annexure 1 & 2)

- Section XX	– Check List for the Tenderers

- Section XXI	– Consignee Detail

8.2 	The relevant details of the required goods/ equipment and services, the terms, conditions and procedure for tendering, tender evaluation, placement of contract, the applicable contract terms and, also, the standard formats to be used for this purpose are incorporated in the above-mentioned documents. The interested tenderers are expected to examine all such details etc to proceed further.

9.	Amendments to TE documents

9.1 	At any time prior to the deadline for submission of tenders, the purchaser may, for any reason deemed fit by it, modify the TE documents by issuing suitable amendment(s) to it.

9.2 	Such an amendment will be notified in writing by registered/speed post or by fax/telex/e-mail, followed by copy of the same by registered post/ courier to all prospective tenderers, which have received the TE documents and will be binding on them. The amendments, if any shall be posted in the websites only.

9.3 	In order to provide reasonable time to the prospective tenderers to take necessary action in preparing their tenders as per the amendment, the purchaser may, at its discretion extend the deadline for the submission of tenders and other allied time frames, which are linked with that deadline.

10.	Clarification of TE documents

  - A tenderer requiring any clarification or elucidation on any issue of the TE documents may take

up the same  with the purchaser in writing. The purchaser will respond in writing to such request

provided the same is received by the purchaser before the pre-bid meeting (unless otherwise

specified in the SIT) .

- C. PREPARATION OF TENDERS

11.	Documents Comprising the Tender

11.1	The Two Stage Tender System, i.e. “Techno – Commercial Tender” and “Price Tender” prepared by the tenderer shall comprise the following:

A)	Techno – Commercial Tender (Un priced Tender)

i)	Earnest money furnished in accordance with GIT clause 19.1 alternatively, documentary evidence as per GIT clause 19.2 for claiming exemption from payment of earnest money.

ii)	Tender Form as per Section X .

iii)	Documentary evidence, as necessary in terms of clauses 5 and 17 establishing that the tenderer is eligible to submit the tender and, also, qualified to perform the contract if its tender is accepted.

iv)	Tenderer/Agent who quotes for goods manufactured by other manufacturer shall furnish Manufacturer’s Authorisation Form as per Section XIV.

v)	Power of Attorney in favour of signatory of TE documents and signatory of Manufacturer’s Authorisation Form

vi)	Documents and relevant details to establish in accordance with GIT clause 18 that the goods and the allied services to be supplied by the tenderer conform to the requirement of the TE documents.

vii)	Performance Statement as per section IX along with relevant copies of orders and end users’ satisfaction certificate.

viii)	Price Schedule(s) as per Section XII filled up with all the details including Make, Model etc. of the goods offered with prices blank (without indicating any prices).

ix)	Certificate of Incorporation in the country of origin.

x)	Checklist as per Section XX.

B)	Price Tender:

The information given at clause no. 11.1 A) ii) & viii) above should be reproduced with the prices indicated.

#### N.B.

      - All pages of the Tender should be page numbered and indexed.

- It is the responsibility of tenderer to go through the TE document to ensure furnishing all required documents in addition to above, if any.

11.2 	The authorized signatory of the tenderer must sign the tender duly stamped at appropriate places and initial all the remaining pages of the tender.

Individuals signing the tender or other documents connected with a contract must specify whether he signs as:

(i) A ‘Sole Proprietor’ of the firm or constituted attorney of such Sole Proprietor.

(ii) A partner of the firm ,if it be a partnership , in which case he must have authority to quote

& to refer to arbitration dispute concerning the business of the partnership either by virtue

of the partnership agreement or a power of attorney;

(iii) Constituted attorney of the firm if it is a company.

N.B. (1)  In case of (ii) above, a copy of the partnership agreement or general power of attorney, in

either ,case, attested by a Notary Public should be furnished, or affidavit on stamped paper

of all the partners admitting execution of the partnership agreement or the general power

of attorney should be furnished.

(2) In case of the partnership firms, where no authority to refer disputes concerning the

business of the partnership has been conferred on any partner, the tender and all other

related documents must be signed by every partner of the firm.

(3) A person signing the tender form or any documents forming part of the contract on behalf

of another shall be deemed to warrantee that he has authority to bind such other persons

and if, on enquiry, it appears that the persons so signing had no authority to do so, the

purchaser may, without prejudice to other civil and criminal remedies, cancel the contract

and hold the signatory liable for all cost and damages;

11.3	A tender, which does not fulfil any of the above requirements and/or gives evasive information/reply against any such requirement, shall be liable to be ignored and rejected.

11.4 	Tender sent by fax/telex/cable/electronically shall be ignored.

12.	Tender currencies

12.1 	The tenderer supplying indigenous goods or already imported goods shall quote only in Indian Rupees.

12.2 	For imported goods if supplied directly from abroad, prices shall be quoted in any freely convertible currency say US Dollar, Euro, GBP or Yen. As regards price(s) for allied services, if any required with the goods, the same shall be quoted in Indian Rupees only if such services are to be performed /undertaken in India. Commission for Indian Agent, if any and if payable shall be indicated in the space provided for in the price schedule and will be payable in Indian Rupees only.

12.3 	Tenders, where prices are quoted in any other way shall be treated as non -responsive and rejected.

13	Tender Prices

13.1 	The Tenderer shall indicate on the Price Schedule provided under Section XI all the specified components of prices shown therein including the unit prices and total tender prices of the goods and services it proposes to supply against the requirement. All the columns shown in the price schedule should be filled up as required. If any column does not apply to a tenderer, same should be clarified as “NA” by the tenderer.

13.2 	If there is more than one schedule in the List of Requirements, the tenderer has the option to submit its quotation for any one or more schedules and, also, to offer special discount for combined schedules. However, while quoting for a schedule, the tenderer shall quote for the complete requirement of goods and services as specified in that particular schedule.

13.3 	The quoted prices for goods offered from within India and that for goods offered from abroad are to be indicated separately in the applicable Price Schedules attached under Section XI.

13.3.1 The price quoted by the tenderer for indigenous goods shall not be higher than the lowest price charged for the goods of the same nature, class or description to an individual/firm/organisation or department of Govt. of India or any state Governments.

13.3.2 For imported goods, the price quoted shall not be higher than the lowest price charged by the tenderer for the goods of the same nature, class or description to a purchaser, domestic or foreign or to any organisation or department of Govt. of India or any state Governments.

13.3.3  If it is found at any stage that the goods as stated have been supplied at a lower price, then that price, with due allowance for elapsed time will be applicable to the present case and the difference in cost would be refunded by the supplier to the purchaser, if the contract has already been concluded.

13.4 	While filling up the columns of the Price Schedule, the following aspects should be noted for compliance:

13.4.1 	For domestic goods or goods of foreign origin located within India, the prices in the corresponding price schedule shall be entered separately in the following manner:

a)	the price of the goods, quoted ex-factory/ ex-showroom/ ex-warehouse/ off-the-shelf, as applicable, including all taxes and duties like sales tax, CST VAT, CENVAT, Custom Duty, Excise Duty etc. already paid or payable or on the previously imported goods of foreign origin quoted ex-showroom etc;

b)	any sales or other taxes and any duties including excise duty, which will be payable on the finished goods in India if the contract is awarded;

c)	charges towards Packing & Forwarding, Inland Transportation, Insurance (local transportation and storage) would be borne by the Supplier from ware house to the consignee site for a period including 3 months beyond date of delivery, Loading/Unloading and other local costs incidental to delivery of the goods to their final destination as specified in the List of Requirements and Price Schedule;

d)	the price of Incidental Services, as mentioned in List of Requirements and Price Schedule;

e)	the prices of Turnkey ( if any), as mentioned in List of Requirements, Technical Specification and Price Schedule; and

f)	the price of annual CMC, as mentioned in List of Requirements, Technical Specification and Price Schedule.

13.4.2	For goods offered from abroad, the prices in the corresponding price schedule shall be entered separately in the following manner:

a)	The price of goods quoted FOB/FCA port of shipment, as indicated in the List of Requirements and Price Schedule;

b)	 The amount of freight and insurance and price of goods quoted CIP port of entry in India.

- The price of goods quoted CIP (name port of destination) in India nearest airport ,seaport basis as indicated in the List of Requirements, Price Schedule and Consignee List;

- Wherever applicable the amount of customs duty as percentage of net CIP value.

e)  The charges for Insurance (local transportation and storage) would be extended and borne by the Supplier from port of entry to the consignee site for a period including 3 months beyond date of delivery. Other local costs and Incidental costs, as specified in the List of Requirements and Price Schedule;

f)	The charges for Incidental Services, as in the List of Requirements and Price Schedule;

g)  Transportation charges from CIP Port of entry to consignee’s site .

h)	The prices of Turnkey (if any), as mentioned in List of Requirements, Technical Specification and Price Schedule; and

i)	The price of annual CMC and AMC, as mentioned in List of Requirements, Technical Specification and Price Schedule.

13.5 	Additional information and instruction on Duties and Taxes:

13.5.1	If the Tenderer desires to ask for excise duty, sales tax/ VAT, Service Tax, Works Contract Tax etc. to be paid extra, the same must be specifically stated. In the absence of any such stipulation the price will be taken inclusive of such duties and taxes and no claim for the same will be entertained later.

13.5.2	Excise Duty:

a)	If reimbursement of excise duty is intended as extra over the quoted prices, the supplier must specifically say so also indicating the rate, quantum and nature of the duty applicable. In the absence of any such stipulation it will be presumed that the prices quoted are firm and final and no claim on account of excise duty will be entertained after the opening of tenders.

b)	If a Tenderer chooses to quote a price inclusive of excise duty and also desires to be reimbursed for variation, if any, in the excise duty during the time of supply, the tenderer must clearly mention the same and also indicate the rate and quantum of excise duty included in its price. Failure to indicate all such details in clear terms may result in rejection of that tender.

c)	Subject to sub clauses 13.5.2 (a) & (b) above, any change in excise duty upward/downward as a result of any statutory variation in excise duty taking place within contract terms shall be allowed to the extent of actual quantum of excise duty paid by the supplier. In case of downward revision in excise duty, the actual quantum of reduction of excise duty shall be reimbursed to the purchaser by the supplier. All such adjustments shall include all reliefs, exemptions, rebates, concession etc. if any obtained by the supplier.

13.5.3	Sales Tax:

If a tenderer asks for sales tax/ VAT, Service Tax and Works Contract Tax to be paid extra, the rate and nature of sales tax applicable should be shown separately. The sales tax / VAT, Service Tax and Works Contract Tax will be paid as per the rate at which it is liable to be assessed or has actually been assessed provided the transaction of sale is legally liable to sales tax / VAT, Service Tax and Works Contract Tax and is payable as per the terms of the contract. If any refund of Tax is received at a later date, the Supplier must return the amount forth-with to the purchaser.

13.5.4 	Octroi Duty and Local Duties & Taxes:

Normally, goods to be supplied to government departments against government contracts are exempted from levy of entry tax/town duty, Octroi duty, terminal tax and other levies of local bodies. However, on some occasions, the local bodies (like town body, municipal body etc.) as per their regulations allow such exemptions only on production of certificate to this effect from the concerned government department. Keeping this in view, the supplier shall ensure that the stores to be supplied by the supplier against the contract placed by the purchaser are exempted from levy of any such duty or tax and, wherever necessary, obtain the exemption certificate from the purchaser.

However, if a local body still insists upon payment of such local duties and taxes, the same should be paid by the supplier to the local body to avoid delay in supplies and possible demurrage charges and obtain a receipt for the same. The supplier should forward the receipt obtained for such payment to the purchaser to enable the purchaser reimburse the supplier and take other necessary action in the matter.

13.5.5	Customs Duty:

The supplier will pay the customs duty and clear the goods for transportation to consignee’s site. The applicable - % rates and amount of custom duty and the corresponding Indian custom tariff number should be shown separately in the price schedule. Duty paid by the supplier shall be re-imbursed on submission of documents. Customs duty exemption certificate (CDEC) wherever applicable shall be issued by the consignee.

13.6	For transportation of imported goods offered from abroad, relevant instructions as incorporated under GCC Clause 10 shall be followed.

13.7	For insurance of goods to be supplied, relevant instructions as provided under GCC Clause 11 shall be followed.

13.8	Unless otherwise specifically indicated in this TE document, the terms FCA, FOB, FAS, CIF, CIP, DDP etc. for imported goods offered from abroad, shall be governed by the rules & regulations prescribed in the current edition of INCOTERMS, published by the International Chamber of Commerce, Paris

13.9	The need for indication of all such price components by the tenderers, as required in this clause (viz., GIT clause 13) is for the purpose of comparison of the tenders by the purchaser and will no way restrict the purchaser’s right to award the contract on the selected tenderer on any of the terms offered.

14.	Indian Agent

14.1	If a foreign tenderer has engaged an agent in India in connection with its tender, the foreign tenderer, in addition to indicating Indian agent’s commission, if any, in a manner described under GIT sub clause 12.2 above, shall also furnish the following information:

a)	The complete name and address of the Indian Agent and its permanent income tax account number as allotted by the Indian Income Tax authority.

b)	The details of the services to be rendered by the agent for the subject requirement.

c)	Details of Service outlets in India, nearest to the consignee(s), to render services during Warranty and CMC period.

d)  A copy of agreement between the Agent & their principal detailing the terms & conditions as well as services and after sales services as above to be rendered by the agent and the precise relationship between them and their mutual interest in the business.

e)   Principal/ manufacturers original proforma invoice with the price bid.

f) The enlistment of the Indian Agent with DGS&D under the compulsory Registration Scheme of Ministry of Finance.

15.	Firm Price

15.1	Unless otherwise specified in the SIT, prices quoted by the tenderer shall remain firm and fixed during the currency of the contract and not subject to variation on any account.

15.2	However, as regards taxes and duties, if any, chargeable on the goods and payable, the conditions stipulated in GIT clause 13 will apply.

16.	Alternative Tenders

16.1	Alternative Tenders are not permitted.

16.2	However the Tenderers can quote alternate models meeting the tender specifications of same manufacturer with single EMD.

17 	Documents Establishing Tenderer’s Eligibility and Qualifications

17.1	Pursuant to GIT clause 11, the tenderer shall furnish, as part of its tender, relevant details and documents establishing its eligibility to quote and its qualifications to perform the contract if its tender is accepted.

17.2	The documentary evidence needed to establish the tenderer’s qualifications shall fulfil the following requirements:

a)	In case the tenderer offers to supply goods, which are manufactured by some other firm, the tenderer has been duly authorised by the goods manufacturer to quote for and supply the goods to the purchaser. The tenderer shall submit the manufacturer’s authorization letter to this effect as per the standard form provided under Section XIV in this document.

b)	The tenderer has the required financial, technical and production capability necessary to perform the contract and, further, it meets the qualification criteria incorporated in the Section IX in these documents.

c)	In case the tenderer is not doing business in India, it is duly represented by an agent stationed in India fully equipped and able to carry out the required contractual functions and duties of the supplier including after sale service, maintenance & repair etc. of the goods in question, stocking of spare parts and fast moving components and other obligations, if any, specified in the conditions of contract and/or technical specifications.

d)	In case the tenderer is an Indian agent/authorized representative quoting on behalf of a foreign manufacturer for the restricted item, the Indian agent/authorized representative is already enlisted under the Compulsory Enlistment Scheme of Ministry of Finance, Govt. of India, operated through Directorate General of Supplies & Disposals (DGS&D), New Delhi.

18.	Documents establishing Good’s Conformity to TE document.

18.1 	The tenderer shall provide in its tender the required as well as the relevant documents like technical data, literature, drawings etc. to establish that the goods and services offered in the tender fully conform to the goods and services specified by the purchaser in the TE documents. For this purpose the tenderer shall also provide a clause-by-clause commentary on the technical specifications and other technical details incorporated by the purchaser in the TE documents to establish technical responsiveness of the goods and services offered in its tender.

18.2 	In case there is any variation and/or deviation between the goods & services prescribed by the purchaser and that offered by the tenderer, the tenderer shall list out the same in a chart form without ambiguity and provide the same along with its tender.

18.3	If a tenderer furnishes wrong and/or misguiding data, statement(s) etc. about technical acceptability of the goods and services offered by it, its tender will be liable to be ignored and rejected in addition to other remedies available to the purchaser in this regard.

19.	Earnest Money Deposit (EMD)

19.1	Pursuant to GIT clauses 8.1 and 11.1(d) the tenderer shall furnish along with its tender, earnest money for amount as shown in the List of Requirements. The earnest money is required to protect the purchaser against the risk of the tenderer’s unwarranted conduct as amplified under sub-clause 19.7 below.

19.2	The tenderers who are currently registered and, also, will continue to remain registered during the tender validity period with Directorate General of Supplies & Disposals or with National Small Industries Corporation, New Delhi for the specific goods as per tender enquiry specification shall be eligible for exemption from EMD. In case the tenderer falls in these categories, it should furnish copy of its valid registration details (with DGS&D or NSIC, as the case may be).

19.3	The earnest money shall be denominated in Indian Rupees or equivalent currencies as per GIT clause 12.2. The earnest money shall be furnished in one of the following forms:

- i) Account Payee Demand Draft

- ii) Banker’s cheque and

- iii) Bank Guarantee

19.4	The demand draft or banker’s cheque shall be drawn on any scheduled commercial bank in India or in country of the tenderer, in favour of the “HLL Lifecare Limited” payable at New Delhi. In case of bank guarantee, the same is to be provided from any scheduled commercial bank in India or in country of the tenderer as per the format specified under Section XIII in these documents.

19.5	The earnest money shall be valid for a period of forty-five (45) days beyond the validity period of the tender. As validity period of Tender as per Clause 20 of GIT is 120 days, the EMD shall be valid for 165 days from Techno – Commercial Tender opening date.

19.6	Unsuccessful tenderers’ earnest money will be returned to them without any interest, after expiry of the tender validity period, but not later than thirty days after conclusion of the resultant contract. Successful tenderer’s earnest money will be returned without any interest, after receipt of performance security from that tenderer.

19.7	Earnest Money is required to protect the purchaser against the risk of the Tenderer’s conduct, which would warrant the forfeiture of the EMD. Earnest money of a tenderer will be forfeited, if the tenderer withdraws or amends its tender or impairs or derogates from the tender in any respect within the period of validity of its tender or if it comes to notice that the information/documents furnished in its tender is incorrect, false, misleading or forged without prejudice to other rights of the purchaser. The successful tenderer’s earnest money will be forfeited without prejudice to other rights of Purchaser if it fails to furnish the required performance security within the specified period.

19.8	In the case of Bank Guarantee furnished from banks outside India (i.e. foreign Banks), it should be authenticated and countersigned by any nationalised bank in India by way of back-to-back counter guarantee.

20.	Tender Validity

20.1	If not mentioned otherwise in the SIT, the tenders shall remain valid for acceptance for a period of 120 days (One hundred and twenty days) after the date of tender opening prescribed in the TE document. Any tender valid for a shorter period shall be treated as unresponsive and rejected.

20.2	In exceptional cases, the tenderers may be requested by the purchaser to extend the validity of their tenders up to a specified period. Such request(s) and responses thereto shall be conveyed by surface mail or by fax/ telex/cable followed by surface mail. The tenderers, who agree to extend the tender validity, are to extend the same without any change or modification of their original tender and they are also to extend the validity period of the EMD accordingly. A tenderer, however, may not agree to extend its tender validity without forfeiting its EMD.

20.3	In case the day up to which the tenders are to remain valid falls on/ subsequently declared a holiday or closed day for the purchaser, the tender validity shall automatically be extended up to the next working day.

21.	Signing and Sealing of Tender

21.1 	The tenderers shall submit their tenders as per the instructions contained in GIT Clause 11.

21.2	Unless otherwise mentioned in the SIT, a tenderer shall submit three copies of its tender marking them as “Original”, “Duplicate” and ”Triplicate”. Duplicate & Triplicate tenders may contain all pages including Technical Literature/Catalogues as per in Original tenders.

21.3	The original and other copies of the tender shall either be typed or written in indelible ink and the same shall be signed by the tenderer or by a person(s) who has been duly authorized to bind the tenderer to the contract. The letter of authorization shall be by a written power of attorney, which shall also be furnished along with the tender.

21.4 	All the copies of the tender shall be duly signed at the appropriate places as indicated in the TE documents and all other pages of the tender including printed literature, if any shall be initialled by the same person(s) signing the tender. The tender shall not contain any erasure or overwriting, except as necessary to correct any error made by the tenderer and, if there is any such correction; the same shall be initialled by the person(s) signing the tender.

21.5	The tenderer is to seal the original and each copy of the tender in separate envelopes, duly marking the same as “Original”, “Duplicate”, ”Triplicate” and so on and writing the address of the purchaser and the tender reference number on the envelopes. The sentence “NOT TO BE OPENED” before __________ (The tenderer is to put the date & time of tender opening) are to be written on these envelopes. The inner envelopes are then to be put in a bigger outer envelope, which will also be duly sealed, marked etc. as above. If the outer envelope is not sealed and marked properly as above, the purchaser will not assume any responsibility for its misplacement, premature opening, late opening etc.

21.6	TE document seeks quotation following two Tender System, in two parts. First part will be known as ‘Techno - Commercial Tender’, and the second part ‘Price Tender’ as specified in clause 11 of GIT. Tenderer shall seal ‘Techno - Commercial Tender’ and ‘Price Tender’ separately and covers will be suitably super scribed. Both these sealed covers shall be put in a bigger cover and sealed and procedure prescribed in Paras 21.1 to 21.5 followed.

D. SUBMISSION OF TENDERS

22.	Submission of Tenders

22.1	Unless otherwise specified, the tenderers are to deposit the tenders in the tender box kept for this purpose at HLL Lifecare Limited, (A Government of India Enterprise), Infrastructure Development Division, “Adarsh” TC 6/1781, Vettamukku, Thirumala P.O, Thiruvananthapuram-695006. In case of bulky tender, which can not be put into tender box, the same shall be submitted by the tenderer by hand to DVP(Tech.) or his nominee, HLL Lifecare Limited, Infrastructure Development Division, “Adarsh” TC 6/1781, Vettamukku, Thirumala P.O, Thiruvananthapuram-695006. The officer receiving the tender will give the tenderer an official receipt duly signed with date and time.

22.2	The tenderers must ensure that they deposit their tenders not later than the closing time and date specified for submission of tenders. It is the responsibility of the tenderer to ensure that their Tenders whether sent by post or by courier or by person, are dropped in the Tender Box by the specified clearing date and time. In the event of the specified date for submission of tender falls on / is subsequently declared a holiday or closed day for the purchaser, the tenders will be received up to the appointed time on the next working day.

23.	Late Tender

23.1	A tender, which is received after the specified date and time for receipt of tenders will be treated as “late” tender and will be ignored.

24.	Alteration and Withdrawal of Tender

24.1	The tenderer, after submitting its tender, is permitted to alter / modify its tender so long as such alterations / modifications are received duly signed, sealed and marked like the original tender, within the deadline for submission of tenders. Alterations / modifications to tenders received after the prescribed deadline will not be considered.

24.2	No tender should be withdrawn after the deadline for submission of tender and before expiry of the tender validity period. If a tenderer withdraws the tender during this period, it will result in forfeiture of the earnest money furnished by the tenderer in its tender.

- E. TENDER OPENING

25. 	Opening of Tenders

25.1	The purchaser will open the tenders at the specified date and time and at the specified place as indicated in the NIT.

In case the specified date of tender opening falls on / is subsequently declared a holiday or closed day for the purchaser, the tenders will be opened at the appointed time and place on the next working day.

25.2	Authorized representatives of the tenderers, who have submitted tenders on time may attend the tender opening provided they bring with them letters of authority from the corresponding tenderers.

The tender opening official(s) will prepare a list of the representatives attending the tender opening. The list will contain the representatives’ names & signatures and corresponding tenderers’ names and addresses.

    - Two - Tender system as mentioned in para 21.6 above will be as follows. The Techno - Commercial Tenders are to be opened in the first instance, at the prescribed time and date as indicated in NIT. During the Techno - Commercial Tender opening, the tender opening official(s) will read the salient features of the tenders like brief description of the goods offered, delivery period, Earnest Money Deposit and any other special features of the tenders, as deemed fit by the tender opening official(s). These Tenders shall be scrutinized and evaluated by the competent committee/ authority with reference to parameters prescribed in the TE document. Thereafter, in the second stage, the Price Tenders of only the Techno - Commercially acceptable offers (as decided in the first stage) shall be opened for further scrutiny and evaluation on a date notified after the evaluation of the Techno – Commercial tender. The prices, special discount if any of the goods offered etc., as deemed fit by tender opening official(s) will be read out.

- F. SCRUTINY AND EVALUATION OF TENDERS

26.	Basic Principle

26.1	Tenders will be evaluated on the basis of the terms & conditions already incorporated in the TE document, based on which tenders have been received and the terms, conditions etc. mentioned by the tenderers in their tenders. No new condition will be brought in while scrutinizing and evaluating the tenders.

27.	Preliminary Scrutiny of Tenders

27.1	The Purchaser will examine the Tenders to determine whether they are complete, whether any computational errors have been made, whether required sureties have been furnished, whether the documents have been properly signed stamped and whether the Tenders are generally in order.

27.2	Prior to the detailed evaluation of Price Tenders, pursuant to GIT Clause 34, the Purchaser will determine the responsiveness of each Tender to the TE Document. For purposes of these clauses, a  responsive Tender is one, which conforms to the technical specifications and all the terms and conditions of the TE Documents without material deviations. Deviations from, or objections or reservations to critical provisions such as those concerning Performance Security , EMD , Tender validity,terms of delivery, liquidated damage,terms of payment,warranty period will be deemed to be a material deviation. The Purchaser’s determination of a Tender’s responsiveness is to be based on the contents of the tender itself without recourse to extrinsic evidence.

27.3	The tenders will be scrutinized to determine whether they are complete and meet the essential and important requirements, conditions etc. as prescribed in the TE document. The tenders, which do not the meet the basic requirements, are liable to be treated as non – responsive and will be summarily ignored.

27.4	The following are some of the important aspects, for which a tender shall be declared non – responsive and will be summarily ignored;

(i)	Tender form as per Section X (signed and stamped) not enclosed

(ii)	Tender is unsigned.

(iii)	Tender validity is shorter than the required period.

(iv)	Required EMD (Amount, validity etc.)/ exemption documents have not been provided.

(v)	Tenderer has quoted for goods manufactured by other manufacturer(s) without the required Manufacturer’s Authorisation Form as per Section XIV.

(vi)	Tenderer has not agreed to give the required performance security.

(vii)	Goods offered are not meeting the tender enquiry Technical specification.

(viii)	Tenderer has not agreed to other essential condition(s) specially incorporated in the tender enquiry like terms of payment, terms of delivery, liquidated damages clause, warranty period.

(ix)	Poor/ unsatisfactory past performance.

(x)	Tenderers who stand deregistered/banned/blacklisted by any Statutory Authorities as per Govt. rules/procedures.

(xi)	Tenderer is not eligible as per GIT Clauses 5.1 & 17.1.

(xii)	Tenderer has not quoted for the entire quantity in the schedule as specified in the List of Requirements .

28.	Minor Informality/Irregularity/Non-Conformity

28.1	If during the preliminary examination, the purchaser find any minor informality and/or irregularity and/or non-conformity in a tender, the purchaser may waive the same provided it does not constitute any material deviation and financial impact and, also, does not prejudice or affect the ranking order of the tenderers. Wherever necessary, the purchaser will convey its observation on such ‘minor’ issues to the tenderer by registered/speed post etc. asking the tenderer to respond by a specified date. If the tenderer does not reply by the specified date or gives evasive reply without clarifying the point at issue in clear terms, that tender will be liable to be ignored.

28.2    The purchaser may seek clarifications of historical nature from the tenderers which has no bearings on prices.

29	Discrepancies in Prices

29.1	If, in the price structure quoted by a tenderer, there is discrepancy between the unit price and the total price (which is obtained by multiplying the unit price by the quantity), the unit price shall prevail and the total price corrected accordingly, unless the purchaser feels that the tenderer has made a mistake in placing the decimal point in the unit price, in which case the total price as quoted shall prevail over the unit price and the unit price corrected accordingly.

29.2	If there is an error in a total price, which has been worked out through addition and/or subtraction of subtotals, the subtotals shall prevail and the total corrected; and

29.3	If there is a discrepancy between the amount expressed in words and figures, the amount in words shall prevail, subject to sub clause 29.1 and 29.2 above.

29.4	If, as per the judgement of the purchaser, there is any such arithmetical discrepancy in a tender, the same will be suitably conveyed to the tenderer by registered / speed post. If the tenderer does not agree to the observation of the purchaser, the tender is liable to be ignored.

30.	Discrepancy between original and copies of Tender

30.1	In case any discrepancy is observed between the text etc. of the original copy and that in the other copies of the same tender set, the text etc. of the original copy shall prevail. Here also, the purchaser will convey its observation suitably to the tenderer by register / speed post and, if the tenderer does not accept the purchaser’s observation, that tender will be liable to be ignored.

31.	Qualification Criteria

31.1	Tenders of the tenderers, who do not meet the required Qualification Criteria prescribed in Section IX, will be treated as non - responsive and will not be considered further.

32.	Conversion of tender currencies to Indian Rupees

32.1	In case the TE document permits the tenderers to quote their prices in different currencies, all such quoted prices of the responsive tenderers will be converted to a single currency viz., Indian Rupees for the purpose of equitable comparison and evaluation, as per the exchange rates established by the Reserve Bank of India for similar transactions, as on the date of ‘Price Tender’ opening.

33.	Schedule-wise Evaluation

33.1	In case the List of Requirements contains more than one schedule, the responsive tenders will be evaluated and compared separately for each schedule. The tender for a schedule will not be considered if the complete requirements prescribed in that schedule are not included in the tender. However, as already mentioned in GIT sub clause 13.2, the tenderers have the option to quote for any one or more schedules and offer discounts for combined schedules. Such discounts wherever applicable will be taken into account to determine the lowest evaluated cost for the purchaser in deciding the successful tenderer for each schedule, subject to tenderer(s) being responsive.

34.	Comparison of Tenders

  - Unless mentioned otherwise in Section – III – Special Instructions to Tenderers and Section – VI – List of Requirements, the comparison of the responsive tenders shall be carried out based on the terms of delivery as asked for in bid document. The quoted turnkey prices and CMC/AMC prices will also be added for comparison/ranking purpose for evaluation if so indicated in the bid document for the required period mentioned in this TED.

35.	Additional Factors and Parameters for Evaluation and Ranking of Responsive Tenders

35.1	Further to GIT Clause 34 above, the purchaser’s evaluation of a tender will include and take into account the following:

i) In the case of goods manufactured in India or goods of foreign origin already located in India, sales tax & other similar taxes and excise duty & other similar duties, Customs Duties, Service Tax, Works Contract Tax etc which will be contractually payable (to the tenderer), on the goods if a contract is awarded on the tenderer; and

ii) in the case of goods of foreign origin offered from abroad, customs duty and other similar import duties/taxes, which will be contractually payable (to the tenderer) on the goods if the contract is awarded on the tenderer.

35.2	The purchaser’s evaluation of tender will also take into account the additional factors, if any, incorporated in SIT in the manner and to the extent indicated therein.

35.3	The Purchaser reserves the right to give the purchase preference to central public sector undertakings as per the instruction in vogue while evaluating, comparing and ranking the responsive tenders.

36.	Tenderer’s capability to perform the contract

36.1	The purchaser, through the above process of tender scrutiny and tender evaluation will determine to its satisfaction whether the tenderer, whose tender has been determined as the lowest evaluated responsive tender is eligible, qualified and capable in all respects to perform the contract satisfactorily. If, there is more than one schedule in the List of Requirements, then, such determination will be made separately for each schedule.

36.2	The above-mentioned determination will, interalia, take into account the tenderer’s financial, technical and production capabilities for satisfying all the requirements of the purchaser as incorporated in the TE document. Such determination will be based upon scrutiny and examination of all relevant data and details submitted by the tenderer in its tender as well as such other allied information as deemed appropriate by the purchaser.

37.	Contacting the Purchaser

37.1	From the time of submission of tender to the time of awarding the contract, if a tenderer needs to contact the purchaser for any reason relating to this tender enquiry and / or its tender, it should do so only in writing.

37.2	In case a tenderer attempts to influence the purchaser in the purchaser’s decision on scrutiny, comparison & evaluation of tenders and awarding the contract, the tender of the tenderer shall be liable for rejection in addition to appropriate administrative actions being taken against that tenderer, as deemed fit by the purchaser.

G. AWARD OF CONTRACT

38.	Purchaser’s Right to accept any tender and to reject any or all tenders

38.1	The purchaser reserves the right to accept in part or in full any tender or reject any or more tender(s) without assigning any reason or to cancel the tendering process and reject all tenders at any time prior to award of contract, without incurring any liability, whatsoever to the affected tenderer or tenderers.

39.	Award Criteria

39.1	Subject to GIT clause 38 above, the contract will be awarded to the lowest evaluated responsive tenderer decided by the purchaser in terms of GIT Clause 36.

40.	Variation of Quantities at the Time of Award/ Currency of Contract

40.1	At the time of awarding the contract, the purchaser reserves the right to increase or decrease by up to twenty five (25) per cent, the quantity of goods and services mentioned in the schedule (s) in the “List of Requirements” without any change in the unit price and other terms & conditions quoted by the tenderer.

40.2	If the quantity has not been increased at the time of the awarding the contract, the purchaser reserves the right to increase by up to twenty five (25) per cent, the quantity of goods and services mentioned in the contract without any change in the unit price and other terms & conditions mentioned in the contract, during the currency of the contract.

41.	Notification of Award

41.1	Before expiry of the tender validity period, the purchaser will notify the successful tenderer(s) in writing, by registered / speed post or by fax/ telex/cable (to be confirmed by registered / speed post) that its tender for goods & services, which have been selected by the purchaser, has been accepted, also briefly indicating therein the essential details like description, specification and quantity of the goods & services and corresponding prices accepted. The successful tenderer must furnish to the purchaser the required performance security within thirty days from the date of dispatch of this notification, failing which the EMD will forfeited and the award will be cancelled. Relevant details about the performance security have been provided under GCC Clause 5 under Section IV.

41.2	The Notification of Award shall constitute the conclusion of the Contract.

42.	Issue of Contract

42.1	Promptly after notification of award, the Purchaser/Consignee will mail the contract form (as per Section XVI) duly completed and signed, in duplicate, to the successful tenderer by registered / speed post.

42.2	Within twenty one days from the date of the contract, the successful tenderer shall return the original copy of the contract, duly signed and dated, to the Purchaser/Consignee by registered / speed post.

42.3	The Purchaser/Consignee reserves the right to issue the Notification of Award consignee wise.

43.	Non-receipt of Performance Security and Contract by the Purchaser/Consignee

43.1	Failure of the successful tenderer in providing performance security and / or returning contract copy duly signed in terms of GIT clauses 41 and 42 above shall make the tenderer liable for forfeiture of its EMD and, also, for further actions by the Purchaser/Consignee against it as per the clause 24 of GCC – Termination of default.

44.	Return of E M D

44.1	The earnest money of the successful tenderer and the unsuccessful tenderers will be returned to them without any interest, whatsoever, in terms of GIT Clause 19.6.

45.	Publication of Tender Result

45.1	The name and address of the successful tenderer(s) receiving the contract(s) will be mentioned in the notice board/bulletin/web site of the purchaser.

46.	Corrupt or Fraudulent Practices

46.1	It is required by all concerned namely the Consignee/Tenderers/Suppliers etc to observe the highest standard of ethics during the procurement and execution of such contracts. In pursuance of this policy, the Purchaser: -

- defines, for the purposes of this provision, the terms set forth below 
as follows:

- “corrupt practice” means the offering, giving, receiving or soliciting of any thing of value to influence the action of a public official in the procurement process or in contract execution; and

- “fraudulent practice” means a misrepresentation of facts in order to influence a procurement process or the execution of a contract to the detriment of the Purchaser, and includes collusive practice among Tenderers (prior to or after Tender submission) designed to establish Tender prices at artificial non-competitive levels and to deprive the Purchaser of the benefits of free and open competition;

- will reject a proposal for award if it determines that the Tenderer recommended for award has engaged in corrupt or fraudulent practices in competing for the contract in question;

- will declare a firm ineligible, either indefinitely or for a stated period of time, to be awarded a contract by the purchaser if it at any time determines that the firm has engaged in corrupt or fraudulent practices in competing for, or in executing the contract.

SECTION - III

SPECIAL INSTRUCTIONS TO TENDERERS

(SIT)

| Sl. No. | GIT Clause No. | Topic | SIT Provision | Page No. |
| --- | --- | --- | --- | --- |
| A | 1 to 7 | Preamble | No Change | 26 |
| B | 8 to 10 | TE documents | No Change | 26 |
| C | 11 to 21 | Preparation of Tenders | No Change | 26 |
| D | 22 to24 | Submission of Tenders | No Change | 26 |
| E | 25 | Tender Opening | No Change | 26 |
| F | 26 to 37 | Scrutiny and Evaluation of Tenders | No Change | 26 |
| G | 38 to 45 | Award of Contract | No Change | 26 |

SPECIAL INSTRUCTIONS TO TENDERERS

(SIT)

The following Special Instructions to Tenderers will apply for this purchase. These special instructions will modify/substitute/supplement the corresponding General Instructions to Tenderers (GIT) incorporated in Section II. The corresponding GIT clause numbers have also been indicated in the text below:

In case of any conflict between the provision in the GIT and that in the SIT, the provision contained in the SIT shall prevail.

A	Preamble

#### No Change

B	TE documents

#### No Change

C	Preparation of Tenders

#### No Change

D	Submission of Tenders

#### No Change

E	Tender Opening

#### No Change

F	Scrutiny and Evaluation of Tenders

#### No Change

G	Award of Contract

No Change

SECTION - IV

GENERAL CONDITIONS OF CONTRACT (GCC)

TABLE OF CLAUSES

| Sl No. | Topic | Page No. |
| --- | --- | --- |
| 1 | Application | 28 |
| 2 | Use of contract documents and information | 28 |
| 3 | Patent Rights | 28 |
| 4 | Country of Origin | 28 |
| 5 | Performance Security | 28 |
| 6 | Technical Specifications and Standards | 29 |
| 7 | Packing and Marking | 29 |
| 8 | Inspection, Testing and Quality Control | 30 |
| 9 | Terms of Delivery | 31 |
| 10 | Transportation of Goods | 31 |
| 11 | Insurance | 31 |
| 12 | Spare parts | 32 |
| 13 | Incidental services | 32 |
| 14 | Distribution of Dispatch Documents for Clearance/Receipt of Goods | 32 |
| 15 | Warranty | 33 |
| 16 | Assignment | 34 |
| 17 | Sub Contracts | 34 |
| 18 | Modification of contract | 34 |
| 19 | Prices | 35 |
| 20 | Taxes and Duties | 35 |
| 21 | Terms and mode of Payment | 35 |
| 22 | Delivery Schedule | 38 |
| 23 | Liquidated Damages | 39 |
| 24 | Termination for default | 39 |
| 25 | Termination for insolvency | 40 |
| 26 | Force Majeure | 40 |
| 27 | Termination for convenience | 40 |
| 28 | Governing language | 41 |
| 29 | Notices | 41 |
| 30 | Resolution of disputes | 41 |
| 31 | Applicable Law | 41 |
| 32 | Withholding and lien in respect of sums claimed | 41 |
| 33 | General/Miscellaneous Clauses | 42 |

SECTION - IV

GENERAL CONDITIONS OF CONTRACT (GCC)

1.	Application

1.1	The General Conditions of Contract incorporated in this section shall be applicable for this purchase to the extent the same are not superseded by the Special Conditions of Contract prescribed under Section V, List of requirements under Section VI and Technical Specification under Section VII of this document.

2.	Use of contract documents and information

2.1	The supplier shall not, without the purchaser’s prior written consent, disclose the contract or any provision thereof including any specification, drawing, sample or any information furnished by or on behalf of the purchaser in connection therewith, to any person other than the person(s) employed by the supplier in the performance of the contract emanating from this TE document. Further, any such disclosure to any such employed person shall be made in confidence and only so far as necessary for the purposes of such performance for this contract.

2.2	Further, the supplier shall not, without the purchaser’s prior written consent, make use of any document or information mentioned in GCC sub-clause 2.1 above except for the sole purpose of performing this contract.

2.3	Except the contract issued to the supplier, each and every other document mentioned in GCC sub-clause 2.1 above shall remain the property of the purchaser and, if advised by the purchaser, all copies of all such documents shall be returned to the purchaser on completion of the supplier’s performance and obligations under this contract.

3.	Patent Rights

3.1	The supplier shall, at all times, indemnify and keep indemnified the purchaser, free of cost, against all claims which may arise in respect of goods & services to be provided by the supplier under the contract for infringement of any intellectual property rights or any other right protected by patent, registration of designs or trademarks. In the event of any such claim in respect of alleged breach of patent, registered designs, trade marks etc. being made against the purchaser, the purchaser shall notify the supplier of the same and the supplier shall, at his own expenses take care of the same for settlement without any liability to the purchaser.

4.	Country of Origin

4.1	All goods and services to be supplied and provided for the contract shall have the origin in India or in the countries with which the Government of India has trade relations.

4.2	The word “origin” incorporated in this clause means the place from where the goods are mined, cultivated, grown, manufactured, produced or processed or from where the services are arranged.

  - The country of origin may be specified in the Price Schedule

5.	Performance Security

5.1	Within thirty (30) days from date of the issue of notification of award by the Purchaser/Consignee, the supplier, shall furnish performance security to the Purchaser/Consignee for an amount equal to ten percent (10%) of the total value of the contract, valid up to sixty (60) days after the date of completion of all contractual obligations by the supplier, including the warranty obligations, initially valid for a period of minimum 30 months from the date of Notification of Award

5.2	The Performance security shall be denominated in Indian Rupees or in the currency of the contract as detailed below:

a)	It shall be in any one of the forms namely Account Payee Demand Draft or Fixed Deposit Receipt drawn from any Scheduled bank in India or Bank Guarantee issued by a Scheduled bank in India, in the prescribed form as provided in section XV of this document in favour of the Purchaser/Consignee. The validity of the Fixed Deposit receipt or Bank Guarantee will be for a period up to sixty (60) days beyond Warranty Period.

5.3	In the event of any failure /default of the supplier with or with out any quantifiable loss to the government/purchaser including furnishing of consignee wise Bank Guarantee for CMC security as per Proforma in Section XV, the amount of the performance security is liable to be forfeited. The Administration Department may do the needful to cover any failure/default of the supplier with or without any quantifiable loss to the Government/purchaser.

5.4	In the event of any amendment issued to the contract, the supplier shall, within twenty-one (21) days of issue of the amendment, furnish the corresponding amendment to the Performance Security (as necessary), rendering the same valid in all respects in terms of the contract, as amended.

5.5	The supplier shall enter into Annual Comprehensive Maintenance Contract as per the ‘Contract Form – B’ in Section XVI with respective consignees, 3 (three) months prior to the completion of Warranty Period. The CMC will commence from the date of expiry of the Warranty Period.

5.6	Subject to GCC sub – clause 5.3 above, the Purchaser/Consignee will release the Performance Security without any interest to the supplier on completion of the supplier’s all contractual obligations including the warranty obligations & after receipt of Consignee wise bank guarantee for CMC security in favour of Head of the Institute of the consignee as per the format in Section XV.

6.	Technical Specifications and Standards

6.1	The Goods & Services to be provided by the supplier under this contract shall conform to the technical specifications and quality control parameters mentioned in ‘Technical Specification’ and ‘Quality Control Requirements’ under Sections VII and VIII of this document.

7.	Packing and Marking

7.1	The packing for the goods to be provided by the supplier should be strong and durable enough to withstand, without limitation, (the entire journey during transit to their final destination as indicated in the contract), rough handling, extreme weather conditions etc. so that there is no damage, deterioration etc. As and if necessary, the size, weights and volumes of the packing cases shall also take into consideration, the remoteness of the final destination of the goods and availability or otherwise of transport and handling facilities at all points during transit up to final destination as per the contract.

7.2	The quality of packing, the manner of marking within & outside the packages and provision of accompanying documentation shall strictly comply with the requirements as provided in Technical Specifications and in SCC. In case the packing requirements are amended due to issue of any amendment to the contract, the same shall also be taken care of by the supplier accordingly.

7.3	Packing instructions:

Unless otherwise mentioned in the Technical Specification and in SCC, the supplier shall make separate packages for each consignee (in case there is more than one consignee mentioned in the contract) and mark each package on three sides with the following with indelible paint of proper quality:

- a. contract number and date

- b. brief description of goods including quantity

- c. packing list reference number

- d. country of origin of goods

- e. consignee’s name and full address and

- f. supplier’s name and address

8.	Inspection, Testing and Quality Control

8.1	The purchaser and/or its nominated representative(s) will /shall be at consignee site, without any extra cost to the purchaser, inspect and/or test the ordered goods and the related services to confirm their conformity to the contract specifications and other quality control details incorporated in the contract. The purchaser shall inform the supplier in advance, in writing, the purchaser’s programme for such inspection and, also the identity of the officials to be deputed for this purpose. The cost towards the transportation, boarding & lodging will be borne by the purchaser and/or its nominated representative(s).

8.2	The Technical Specification and Quality Control Requirements incorporated in the contract shall specify what inspections and tests are to be carried out and, also, where and how they are to be conducted. If such inspections and tests are conducted in the premises of the supplier or its subcontractor(s), all reasonable facilities and assistance, including access to relevant drawings, design details and production data, shall be furnished by the supplier to the purchaser’s inspector at no charge to the purchaser.

  - If during such inspections and tests the contracted goods fail to conform to the required specifications and standards, the purchaser’s inspector may reject them and the supplier shall either replace the rejected goods or make all alterations necessary to meet the specifications and standards, as required, free of cost to the purchaser and resubmit the same to the purchaser’s inspector for conducting the inspections and tests again.

8.4	In case the contract stipulates pre-despatch inspection of the ordered goods at supplier’s premises, the supplier shall put up the goods for such inspection to the purchaser’s inspector well ahead of the contractual delivery period, so that the purchaser’s inspector is able to complete the inspection within the contractual delivery period. The goods, should, on no account be dispatched /delivered without getting the same inspected and passed by the inspecting officer stipulated in the contract .

8.5	If the supplier tenders the goods to the purchaser’s inspector for inspection at the last moment without providing reasonable time to the inspector for completing the inspection within the contractual delivery period, the inspector may carry out the inspection and complete the formality beyond the contractual delivery period at the risk and expense of the supplier. The fact that the goods have been inspected after the contractual delivery period will not have the effect of keeping the contract alive and this will be without any prejudice to the legal rights and remedies available to the purchaser under the terms & conditions of the contract.

8.6	The purchaser’s/consignee’s contractual right to inspect, test and, if necessary, reject the goods after the goods’ arrival at the final destination shall have no bearing of the fact that the goods have previously been inspected and cleared by purchaser’s inspector during pre-despatch inspection mentioned above. On rejection the supplier shall remove such stores within 14 days of the date of intimation of such rejection from consignee’s premises. If such goods are not removed by the supplier within the period aforementioned, the purchaser/consignee may remove the rejected stores and either return the same to the supplier at his risk and cost by such mode of transport as purchaser/consignee may decide, or dispose of such goods at the supplier’s risk to recover any expense incurred in connection with such disposals and also the cost of the rejected stores if already paid for.

8.7	Goods accepted by the purchaser/consignee and/or its inspector at initial inspection and in final inspection in terms of the contract shall in no way dilute purchaser’s/consignee’s right to reject the same later, if found deficient in terms of the warranty clause of the contract, as incorporated under GCC Clause 15.

8.8	Principal/ Foreign supplier shall also have the equipment inspected by recognised/ reputed agency like SGS, Lloyd or equivalent (acceptable to the purchaser) prior to despatch at the supplier’s cost and furnish necessary certificate from the said agency in support of their claim.

9.	Terms of Delivery

9.1	Goods shall be delivered by the supplier in accordance with the terms of delivery as specified in the list of requirement.

10.	Transportation of Goods

10.1.1	Instructions for transportation of imported goods offered from abroad:

The supplier shall not arrange part-shipments and/or transhipment without the express/prior written consent of the purchaser. The supplier is required under the contract to deliver the goods under DDP consignee site The shipment shall be made by Indian flag vessel or by vessels belonging to the conference lines in which India is a member country through India’s forwarding agents/coordinators. In case the forwarding agent/coordinators are unable to provide timely adequate space in Indian flag vessel or by vessels belonging to the conference lines, the supplier shall arrange shipment through any available vessel to adhere to the delivery schedule given in the contract. The procedure followed should be intimated to the purchaser.

    - In case of airlifting of imported goods offered from abroad, the same will be done only through

the National Carrier i.e. Air India wherever applicable. In case the National Carrier is not

available, any other airlines available for early delivery may be arranged.

10.1.3  The procedure followed as above should be intimated to the purchaser.

The goods will be custom cleared by the supplier/Indian Agent after paying customs duty and will be transported to the consignee’s site as per terms of the contract.

  - Instructions for transportation of domestic goods including goods already imported by the

supplier.

In case no instruction is provided in this regard in the SCC, the supplier will arrange transportation of the ordered goods as per its own procedure. The supplier shall be responsible for all lose ,destructions,  damage or deterioration of or to the goods from any cause whatsoever while the goods after approval by the inspector are awaiting despatch or delivery.

11.	Insurance:

11.1	Unless otherwise instructed in the SCC, the supplier shall make arrangements for insuring the goods at his cost against loss or damage incidental to manufacture or acquisition, transportation, storage and delivery in the following manner:

- in case of supply of domestic goods on  Consignee site basis, the supplier shall be responsible till the entire stores contracted for arrival in good condition at destination. The transit risk in this respect shall be covered by the Supplier by getting the stores duly insured for an amount equal to 110% of the value of the goods from “warehouse to warehouse” (consignee site) on all risk basis. The insurance cover shall be obtained by the Supplier and should be valid till 3 months after the receipt of goods by the Consignee.

ii)	supply of the imported goods DDP basis, the supplier shall arrange and pay for marine/air insurance making the consignee as beneficiary. The additional extended Insurance (local transportation and storage) would also be borne by the Supplier from the port of entry to the consignee site for a period including 3 months beyond date of delivery. The insurance shall be taken for an amount equal to 110% of overall expenditure to be incurred by the purchaser for receiving the goods at consignee’s site.

If the equipment is not commissioned and handed over to the consignee within 3 months, the insurance will be got extended by the supplier at their cost till the successful installation, testing, commissioning and handing over of the goods to the consignee is completed. In case the delay in the installation and commissioning is due to handing over of the site to the supplier by the consignee, such extensions of the insurance will still be done by the supplier, but the insurance extension charges at actuals will be reimbursed.

12.	Spare parts

  - If specified in the List of Requirements and in the resultant contract, the supplier shall supply/provide any or all of the following materials, information etc. pertaining to spare parts manufactured and/or supplied by the supplier including their prices:

a)         Spare Parts list and prices of parts, consumables should be mentioned clearly and quoted.

Bidder should also mention regarding the availably of spares for at least eight years.

b)	The spare parts as selected by the Purchaser/Consignee to be purchased from the supplier, subject to the condition that such purchase of the spare parts shall not relieve the supplier of any contractual obligation including warranty obligations; and

c)	In case the production of the spare parts is discontinued:

i)	Sufficient advance notice to the Purchaser/Consignee before such discontinuation to provide adequate time to the purchaser to purchase the required spare parts etc., and

ii)	Immediately following such discontinuation, providing the Purchaser/Consignee, free of cost, the designs, drawings, layouts and specifications of the spare parts, as and if requested by the Purchaser/Consignee.

12.2	Supplier shall carry sufficient inventories to assure ex-stock supply of consumable spares for the goods so that the same are supplied to the Purchaser/Consignee promptly on receipt of order from the Purchaser/Consignee.

13.	Incidental services

13.1	Subject to the stipulation, if any, in the SCC (Section – V), List of Requirements (Section – VI) and the Technical Specification (Section – VII), the supplier shall be required to perform the following services.

i)	Installation & commissioning, Supervision and Demonstration of the goods

ii)	Providing required jigs and tools for assembly, minor civil works required for the completion of the installation.

iii)	Training of Consignee’s Staff, operators etc. for operating and maintaining the goods

iv)	Supplying required number of operation & maintenance manual for the goods

14.	Distribution of Dispatch Documents for Clearance/Receipt of Goods

The supplier shall send all the relevant despatch documents well in time to the Purchaser/Consignee to enable the Purchaser/Consignee clear or receive (as the case may be) the goods in terms of the contract.

Unless otherwise specified in the SCC, the usual documents involved and the drill to be followed in general for this purpose are as follows.

A) 	For Domestic Goods, including goods already imported by the supplier under its own arrangement

Within 24 hours of despatch, the supplier shall notify the purchaser, consignee, and others concerned if mentioned in the contract, the complete details of despatch and also supply the following documents to them by registered post / speed post (or as instructed in the contract):

- Four copies of supplier’s invoice showing contract number, goods description, quantity, unit price and total amount;

- Two copies of packing list identifying contents of each package;

- Inspection certificate issued by the nominated Inspection agency, if any.

- Certificate of origin;

- Insurance Certificate as per GCC Clause 11.

- Manufacturers/Supplier’s warranty certificate & In-house inspection certificate.

B)	For goods imported from abroad

Within 24 hours of despatch, the supplier shall notify the purchaser, consignee, and others concerned if mentioned in the contract, the complete details of despatch and also supply the following documents to them by registered post / speed post (or as instructed in the contract). In case of CIP contracts where custom clearance is to be done by purchaser any delay or demurrage occurred during the customs clearance on account of the non-availability of technical support/ clarifications /documents from the supplier shall be borne by the supplier:

- Four copies of supplier’s invoice showing contract number, goods description, quantity, unit price and total amount;

- Original and four copies of the negotiable clean, on-board Bill of Lading/Airway bill, marked freight pre paid and four copies of non-negotiable Bill of Lading/Airway bill;

- Four Copies of packing list identifying contents of each package;

- Insurance Certificate as per GCC Clause 11.

- Manufacturer’s/Supplier’s warranty certificate;

(vi)	Inspection Certificate for the despatched equipments issued by   recognized/ reputed    agency like SGS, Lloyd or equivalent (acceptable to the purchaser) prior to despatch

(vii)	Manufacturer’s own factory inspection report;

- Certificate of origin

- Port of Loading;

- Port of Discharge and

- Expected date of arrival.

15.	Warranty

15.1	The supplier warrants comprehensively that the goods supplied under the contract is new, unused and incorporate all recent improvements in design and materials unless prescribed otherwise by the purchaser in the contract. The supplier further warrants that the goods supplied under the contract shall have no defect arising from design, materials (except when the design adopted and / or the material used are as per the Purchaser’s/Consignee’s specifications) or workmanship or from any act or omission of the supplier, that may develop under normal use of the supplied goods under the conditions prevailing in India.

15.2	This warranty shall remain valid for the period as mentioned in the list of requirement after the goods or any portion thereof as the case may be, have been delivered to the final destination and installed and commissioned at the final destination and accepted by the Purchaser/Consignee in terms of the contract, unless specified otherwise in the SCC.

  - No conditional warranty will be acceptable.

  - Warranty as well as Comprehensive Maintenance contract will be inclusive of all accessories and Turnkey work.

  - Replacement and repair will be under taken for the defective goods.

  - Proper marking has to be made for all spares for identification like printing of installation and repair dates.

15.3	In case of any claim arising out of this warranty, the Purchaser/Consignee shall promptly notify the same in writing to the supplier. The period of the warranty will be as per G.C.C clause number 15.2 above irrespective of any other period mentioned elsewhere in the bidding documents.

15.4	Upon receipt of such notice, the supplier shall, within 24 hours on a 24(hrs) X 7 (days) X 365 (days) basis respond to take action to repair or replace the defective goods or parts thereof, free of cost, at the ultimate destination. The supplier shall take over the replaced parts/goods after providing their replacements and no claim, whatsoever shall lie on the purchaser for such replaced parts/goods thereafter. The penalty clause for non rectification will be applicable as per tender conditions

15.5	In the event of any rectification of a defect or replacement of any defective goods during the warranty period, the warranty for the rectified/replaced goods shall be extended to a further period of twenty four (24) months from the date such rectified / replaced goods starts functioning to the satisfaction of the purchaser. In case the supplier is not able to rectify the defects to the full satisfaction of the purchaser the goods shall have to be replaced with a new one and fresh warranty as per Clause 15.2 above shall be applicable. The decision of the purchaser in this respect shall be final and binding on the supplier .

15.6	If the supplier, having been notified, fails to respond to take action to repair or replace the defect(s) within 24 hours on a 24(hrs) X 7 (days) X 365 (days) basis, the purchaser may proceed to take such remedial action(s) as deemed fit by the purchaser, at the risk and expense of the supplier and without prejudice to other contractual rights and remedies, which the purchaser may have against the supplier.

15.7	During Warranty period, the supplier is required to visit at each consignee’s site at least once in 6 months commencing from the date of the installation for preventive maintenance of the goods

15.8	The Purchaser/Consignee reserve the rights to enter into Annual Comprehensive Maintenance Contract between Consignee and the Supplier for the period as mentioned in Section VII, Technical Specifications after the completion of warranty period.

15.9	The supplier along with its Indian Agent and the CMC provider shall ensure continued supply of the spare parts for the machines and equipments supplied by them to the purchaser for 10 years from the date of installation and handing over.

16.	Assignment

16.1	The Supplier shall not assign, either in whole or in part, its contractual duties, responsibilities and obligations to perform the contract, except with the Purchaser’s prior written permission.

17.	Sub Contracts

17.1	The Supplier shall notify the Purchaser in writing of all sub contracts awarded under the contract if not already specified in its tender. Such notification, in its original tender or later, shall not relieve the Supplier from any of its liability or obligation under the terms and conditions of the contract.

17.2	Sub contract shall be only for bought out items and sub-assemblies.

17.3	Sub contracts shall also comply with the provisions of GCC Clause 4 (“Country of Origin”).

18.	Modification of contract

18.1	If necessary, the purchaser may, by a written order given to the supplier at any time during the currency of the contract, amend the contract by making alterations and modifications within the general scope of contract in any one or more of the following:

a) Specifications, drawings, designs etc. where goods to be supplied under the contract are to be specially manufactured for the purchaser,

b) Mode of packing,

c) Incidental services to be provided by the supplier

d) Mode of despatch,

e) Place of delivery, and

f) Any other area(s) of the contract, as felt necessary by the purchaser depending on the merits of the case.

18.2	In the event of any such modification/alteration causing increase or decrease in the cost of goods and services to be supplied and provided, or in the time required by the supplier to perform any obligation under the contract, an equitable adjustment shall be made in the contract price and/or contract delivery schedule, as the case may be, and the contract amended accordingly. If the supplier doesn’t agree to the adjustment made by the Purchaser/Consignee, the supplier shall convey its views to the Purchaser/Consignee within twenty-one days from the date of the supplier’s receipt of the Purchaser’s/Consignee’s amendment / modification of the contract.

19.	Prices

19.1	Prices to be charged by the supplier for supply of goods and provision of services in terms of the contract shall not vary from the corresponding prices quoted by the supplier in its tender and incorporated in the contract except for any price adjustment authorised in the SCC.

20.	Taxes and Duties

20.1	Supplier shall be entirely responsible for all taxes, duties, fees, levies etc. incurred until delivery of the contracted goods to the purchaser.

  - Further instruction, if any, shall be as provided in the SCC.

  - No form ‘D’ to be issued for concessional CST as the same is no longer applicable.

21.	Terms and Mode of Payment

## 21.1	Payment Terms

Payment shall be made subject to recoveries, if any, by way of liquidated damages or any other charges as per terms & conditions of contract in the following manner.

A)	Payment for Domestic Goods Or Foreign Origin Located Within India.

Payment shall be made in Indian Rupees as specified in the contract in the following manner:

- On delivery:

80 % payment of the contract price shall be paid on receipt of goods in good condition and upon the submission of the following documents:

- Four copies of supplier’s invoice showing contract number, goods description, quantity, unit price and total amount;

- Consignee Receipt Certificate as per Section XVII in original issued by the authorized representative of the consignee;

- Two copies of packing list identifying contents of each package;

- Inspection certificate issued by the nominated Inspection agency, if any.

- Insurance Certificate as per GCC Clause 11

- Certificate of origin.

- On Acceptance:

Balance 20 % payment would be made against ‘Final Acceptance Certificate’ as per Section XVIII of goods to be issued by the consignees subject to recoveries, if any, either on account of non-rectification of defects/deficiencies not attended by the Supplier or otherwise.

B)	Payment for Imported Goods:

Payment for foreign currency portion shall be made in the currency as specified in the contract in the following manner:

For contracts placed on DDP (consignee site ) basis

    - On delivery:

Eighty (80) % of the net CIP price (CIP price less Indian Agency commission ) of the goods shipped shall be paid through irrevocable, non-transferable Letter of Credit (LC) opened in favour of the supplier in a bank in his country and upon submission of documents specified hereunder:

      - Four copies of supplier’s invoice showing contract number, goods description, quantity, unit price and total amount.

      - Original and four copies of the negotiable clean, on–board Bill of Lading/Airway bill, marked freight pre paid and four copies of non-negotiable Bill of Lading/Airway bill;

      - Four copies of packing list identifying contents of each package;

      - Insurance Certificate as per GCC Clause 11

      - Manufacturer’s/Supplier’s warranty certificate;

      - Manufacturer’s own factory inspection report and

      - Certificate of origin by the chamber of commerce of the concerned country,

      - Inspection Certificate for the despatched equipments issued by recognized/reputed agency like SGS, Lloyd or equivalent (acceptable to the purchaser) prior to despatch.

      - Consignee Receipt Certificate as per Section XVII in original issued by the authorized representative of the consignee.

b)   On Acceptance:

Balance payment of 20% of net CIP price of goods would be made against ‘Final Acceptance Certificate’ as per Section XVIII to be issued by the consignees through irrevocable, non- transferable Letter of  Credit (LC) opened in favour of the Foreign Principal in a bank in his country, Subject to recoveries, if any.

c)	For contracts on DDP basis

Payment of custom duty amount with Custom Duty Exemption Certificate (CDEC), if applicable, customs clearance and handling charges, loading/ unloading, inland transportation, incidental costs till consignee site & incidental services (including installation & commissioning, supervision, demonstration and training) will be paid in Indian Rupees to the Indian agent at actual not exceeding the quoted rates on proof of 100% payment to the foreign principal.

d)	Payment of Indian Agency Commission:

Indian Agency commission will be paid to the manufacturer’s agent in the local currency for an amount in Indian rupees indicated in the relevant Price Schedule (as per prevailing rate of exchange ruling on the date of Contract) and shall not be subject to further escalation / exchange variation. Payment shall be paid in Indian Rupees to the Indian Agent on proof of 100 % payment to the Foreign Principal.

C)	Payment of Turnkey, if any:

Turnkey payment will be made to the manufacturer’s agent in Indian rupees indicated in the relevant Price Schedule (as per prevailing rate of exchange ruling on the date of Contract) and shall not be subject to further escalation / exchange variation. Payment shall be made in Indian Rupees to the Indian Agent on proof of 100 % payment to the Foreign Principal.

D)	Payment for Annual Comprehensive Maintenance Contract Charges:

The consignee will enter into CMC with the supplier at the rates as stipulated in the contract. The payment of CMC will be made on six monthly basis after satisfactory completion of said period, duly certified by the consignee on receipt of bank guarantee for an amount equivalent to 2.5 % of the cost of the equipment as per contract in the prescribed format given in Section XV valid till 2 months after expiry of entire CMC period.

21.2	The supplier shall not claim any interest on payments under the contract.

21.3	Where there is a statutory requirement for tax deduction at source, such deduction towards income tax and other tax as applicable will be made from the bills payable to the Supplier at rates as notified from time to time.

21.4	Irrevocable & non – transferable LC shall be opened by the purchaser. However, if the supplier requests specifically to open confirmed LC, the extra charges would be borne by the supplier. If LC is required to be extended and/or amended for reasons not attributable to the purchaser/consignee, the charges thereof shall be borne by the supplier.

21.5	The payment shall be made in the currency / currencies authorised in the contract.

21.6	The supplier shall send its claim for payment in writing, when contractually due, along with relevant documents etc., duly signed with date, to respective consignees.

21.7	While claiming payment, the supplier is also to certify in the bill that the payment being claimed is strictly in terms of the contract and all the obligations on the part of the supplier for claiming that payment has been fulfilled as required under the contract.

21.8	While claiming reimbursement of duties, taxes etc. (like sales tax, excise duty, custom duty) from the Purchaser/Consignee, as and if permitted under the contract, the supplier shall also certify that, in case it gets any refund out of such taxes and duties from the concerned authorities at a later date, it (the supplier) shall refund to the Purchaser/Consignee forthwith.

21.9	In case where the supplier is not in a position to submit its bill for the balance payment for want of receipted copies of Inspection Note from the consignee and the consignee has not complained about the non-receipt, shortage, or defects in the supplies made, balance amount will be paid by the paying authority without consignee’s receipt certificate after three months from the date of the preceding part payment for the goods in question, subject to the following conditions:

(a)	The supplier will make good any defect or deficiency that the consignee (s) may report within six months from the date of despatch of goods.

(b)	Delay in supplies, if any, has been regularized.

(c)	The contract price where it is subject to variation has been finalized.

(d)	The supplier furnishes the following undertakings:

“I/We, __________ certify that I/We have not received back the Inspection Note duly receipted by the consignee or any communication from the purchaser or the consignee about non-receipt, shortage or defects in the goods supplied. I/We ______ agree to make good any defect or deficiency that the consignee may report within three months from the date of receipt of this balance payment.

22.	Delivery Schedule

22.1	The supplier shall deliver of the goods and perform the services under the contract within the time schedule specified by the Purchaser/Consignee in the List of Requirements and as incorporated in the contract. The time for and the date of delivery of the goods stipulated in the schedule shall be deemed to be of the essence of contract and the delivery must be completed not later than the date(s) as specified in the Contract.

22.2	Subject to the provision under GCC clause 26, any unexcused delay by the supplier in maintaining its contractual obligations towards delivery of goods and performance of services shall render the supplier liable to any or all of the following sanctions:

(i) imposition of liquidated damages,

(ii) forfeiture of its performance security and

(iii) termination of the contract for default.

22.3	If at any time during the currency of the contract, the supplier encounters conditions hindering timely delivery of the goods and performance of services, the supplier shall promptly inform the Purchaser/Consignee in writing about the same and its likely duration and make a request to the Purchaser/Consignee for extension of the delivery schedule accordingly. On receiving the supplier’s communication, the Purchaser/Consignee shall examine the situation as soon as possible and, at its discretion, may agree to extend the delivery schedule, with or without liquidated damages for completion of supplier’s contractual obligations by issuing an amendment to the contract.

22.4	When the period of delivery is extended due to unexcused delay by the supplier, the amendment letter extending the delivery period shall, interalia contain the following conditions:

- (a) The Purchaser/Consignee shall recover from the supplier, under the provisions of the clause 23 of the General Conditions of Contract, liquidated damages on the goods and services, which the Supplier has failed to deliver within the delivery period stipulated in the contract.

- (b) That no increase in price on account of any ground, whatsoever, including any stipulation in the contract for increase in price on any other ground and, also including statutory increase in or fresh imposition of customs duty, excise duty, sales tax/ VAT, Service Tax and Works Contract Tax or on account of any other tax or duty which may be levied in respect of the goods and services specified in the contract, which takes place after the date of delivery stipulated in the contract shall be admissible on such of the said goods and services as are delivered and performed after the date of the delivery stipulated in the contract.

- (c) But nevertheless, the Purchaser/Consignee shall be entitled to the benefit of any decrease in price on account of reduction in or remission of customs duty, excise duty, sales tax/ VAT, Service Tax and Works Contract Tax or any other duty or tax or levy or on account of any other grounds, which takes place after the expiry of the date of delivery stipulated in the contract.

  - The supplier shall not dispatch the goods after expiry of the delivery period. The supplier is required to apply to the Purchaser/Consignee for extension of delivery period and obtain the same before despatch. In case the supplier dispatches the goods without obtaining an extension, it would be doing so at its own risk and no claim for payment for such supply and / or any other expense related to such supply shall lie against the purchaser.

  - Passing of Property:

    - The property in the goods shall not pass to the purchaser unless and until the goods have

been delivered to the consignee in accordance with the conditions of the contract.

    - Where there is a contract for sale of specific goods and the supplier is bound to do something to the goods for the purpose of putting them into a   deliverable state the property does not pass  until such thing is done.

    - Unless otherwise agreed, the goods remain at the supplier’s risk until the property therein is transferred to the purchaser.

23.	Liquidated damages

23.1	Subject to GCC clause 26, if the supplier fails to deliver any or all of the goods or fails to perform the services within the time frame(s) incorporated in the contract, the Purchaser/Consignee  shall, without prejudice to other rights and remedies available to the Purchaser/Consignee  under the contract, deduct from the contract price, as liquidated damages, a sum equivalent to 0.5% per week of delay or part thereof on delayed supply of goods and/or services until actual delivery or performance subject to a maximum of 10% of the contract price. Once the maximum is reached Purchaser/Consignee may consider termination of the contract as per GCC 24.

During the above-mentioned delayed period of supply and / or performance, the conditions incorporated under GCC sub-clause 22.4 above shall also apply.

24.	Termination for default

24.1	The Purchaser/Consignee , without prejudice to any other contractual rights and remedies available to it (the Purchaser/Consignee ), may, by written notice of default sent to the supplier, terminate the contract in whole or in part, if the supplier fails to deliver any or all of the goods or fails to perform any other contractual obligation(s) within the time period specified in the contract, or within any extension thereof granted by the Purchaser/Consignee  pursuant to GCC sub-clauses 22.3 and 22.4.

24.2	In the event of the Purchaser/Consignee terminates the contract in whole or in part, pursuant to GCC sub-clause 24.1 above, the Purchaser/Consignee will forfeit the performance security and may procure goods and/or services similar to those cancelled, with such terms and conditions and in such manner as it deems fit . The supplier shall be liable to the Purchaser/Consignee for the extra expenditure, if any, incurred by the Purchaser/Consignee for arranging such procurement.

24.3	Unless otherwise instructed by the Purchaser/Consignee, the supplier shall continue to perform the contract to the extent not terminated.

25.	Termination for insolvency

25.1	If the supplier becomes bankrupt or otherwise insolvent, the purchaser reserves the right to terminate the contract at any time, by serving written notice to the supplier without any compensation, whatsoever, to the supplier, subject to further condition that such termination will not prejudice or affect the rights and remedies which have accrued and / or will accrue thereafter to the Purchaser/Consignee.

26.	Force Majeure

26.1	Notwithstanding the provisions contained in GCC clauses 22, 23 and 24, the supplier shall not be liable for imposition of any such sanction so long the delay and/or failure of the supplier in fulfilling its obligations under the contract is the result of an event of Force Majeure.

26.2	For purposes of this clause, Force Majeure means an event beyond the control of the supplier and not involving the supplier’s fault or negligence and which is not foreseeable and not brought about at the instance of , the party claiming to be affected by such event and which has caused the non – performance or delay in performance. Such events may include, but are not restricted to, acts of the Purchaser/Consignee  either in its sovereign or contractual capacity, wars or revolutions, hostility, acts of public enemy, civil commotion, sabotage, fires, floods, explosions, epidemics, quarantine restrictions, strikes excluding by its employees , lockouts excluding  by its management, and freight embargoes.

26.3	If a Force Majeure situation arises, the supplier shall promptly notify the Purchaser/Consignee in writing of such conditions and the cause thereof within twenty one days of occurrence of such event. Unless otherwise directed by the Purchaser/Consignee in writing, the supplier shall continue to perform its obligations under the contract as far as reasonably practical, and shall seek all reasonable alternative means for performance not prevented by the Force Majeure event.

26.4	If the performance in whole or in part or any obligation under this contract is prevented or delayed by any reason of Force Majeure for a period exceeding sixty days, either party may at its option terminate the contract without any financial repercussion on either side.

26.5	In case due to a Force Majeure event the Purchaser/Consignee is unable to fulfil its contractual commitment and responsibility, the Purchaser/Consignee will notify the supplier accordingly and subsequent actions taken on similar lines described in above sub-paragraphs.

27.	Termination for convenience

27.1	The Purchaser/Consignee  reserves the right to terminate the contract, in whole or in part for its (Purchaser’s/Consignee ’s) convenience, by serving written notice on the supplier at any time during the currency of the contract. The notice shall specify that the termination is for the convenience of the Purchaser/Consignee. The notice shall also indicate interalia, the extent to which the supplier’s performance under the contract is terminated, and the date with effect from which such termination will become effective.

27.2	The goods and services which are complete and ready in terms of the contract for delivery and performance within thirty days after the supplier’s receipt of the notice of termination shall be accepted by the Purchaser/Consignee following the contract terms, conditions and prices. For the remaining goods and services, the Purchaser/Consignee may decide:

a)	To get any portion of the balance completed and delivered at the contract terms, conditions and prices; and / or

b)	To cancel the remaining portion of the goods and services and compensate the supplier by paying an agreed amount for the cost incurred by the supplier towards the remaining portion of the goods and services.

28.	Governing language

28.1	The contract shall be written in English language following the provision as contained in GIT clause 4. All correspondence and other documents pertaining to the contract, which the parties exchange, shall also be written accordingly in that language.

29.	Notices

29.1	Notice, if any, relating to the contract given by one party to the other, shall be sent in writing or by cable or telex or facsimile and confirmed in writing. The procedure will also provide the sender of the notice, the proof of receipt of the notice by the receiver. The addresses of the parties for exchanging such notices will be the addresses as incorporated in the contract.

29.2	The effective date of a notice shall be either the date when delivered to the recipient or the effective date specifically mentioned in the notice, whichever is later.

30.	Resolution of disputes

30.1	If dispute or difference of any kind shall arise between the Purchaser/Consignee and the supplier in connection with or relating to the contract, the parties shall make every effort to resolve the same amicably by mutual consultations.

30.2	If the parties fail to resolve their disputes of differences by such mutual consultation within twenty-one days of its occurrence, then, unless otherwise provided in the SCC, either the Purchaser/Consignee or the supplier may give notice to the other party of its intention to commence arbitration, as hereinafter provided, the applicable arbitration procedure will be as per the Arbitration and Conciliation Act,1996 of India. Such dispute or difference shall be referred to the sole arbitration of an officer in the Ministry of Law and Justice, appointed to be the arbitrator by the CMD of HLL. The award of the arbitrator shall be final and binding on the parties to the contract subject to the provision that the Arbitrator shall give reasoned award in case the value of claim in reference exceeds Rupees One lakhs (Rs. 1,00,000/-)

30.3	Venue: The venue of arbitration shall be Delhi/New Delhi( India).

31.	Applicable Law

The contract shall be governed by and interpreted in accordance with the laws of India for the time being in force.

32.      Withholding and Lien in respect of sums claimed

Whenever any claim for payment arises under the contract against the supplier the purchaser shall

be entitled to withhold and also have a lien to retain such sum from the security deposit or sum of

money arising out of under any other contract made by the supplier with the purchaser, pending

finalization or adjudication of any such claim.

It is an agreed term of the contract that the sum of money so withheld or retained under

the lien referred to above ,by the purchaser, will be kept withheld or retained till the claim arising

about of or under the contract is determined by the Arbitrator or by the competent court as the

case may be ,and the supplier will have no claim for interest or damages whatsoever on any

account in respect of such withholding or retention.

33.	General/ Miscellaneous Clauses

33.1    Each member/constituent of the Supplier/its Indian Agent/CMC/AMC Provider, in case of

consortium shall be jointly and severally liable to and responsible for all obligations towards the

Purchaser/Consignee for performance of contract/services including that of its

Associates/Sub Contractors under the Contract.

  - The Supplier/its Indian Agent/CMC/AMC Provider shall at all times, indemnify and keep

indemnified the Purchaser against all claims/damages etc. for any

infringement of any Intellectual Property Rights (IPR) while providing its services under

CMC/AMC or the Contract.

  - The Supplier/its Agent/CMC/AMC Provider shall, at all times, indemnify and keep indemnified the Purchaser/Consignee against any claims in respect of any damages or compensation payable in consequences of any accident or injury sustained or suffered by its employees or agents or by any other third party resulting from or by any action, omission or operation conducted by or on behalf of the supplier/its associate/affiliate etc.

  - All claims regarding indemnity shall survive the termination or expiry of the contract .

SECTION – V

SPECIAL CONDITIONS OF CONTRACT (SCC)

The following Special Conditions of Contract (SCC) will apply for this purchase. The corresponding clauses of General Conditions of Contract (GCC) relating to the SCC stipulations have also been incorporated below.

These Special Conditions will modify/substitute/supplement the corresponding (GCC) clauses.

Whenever there is any conflict between the provision in the GCC and that in the SCC, the provision contained in the SCC shall prevail.

  - Nil

SECTION - VI

LIST OF REQUIREMENTS

Part I

| Sch. No. | Equipment | Total Quantity |
| --- | --- | --- |
| 1 | Dual Head SPECT (CT) Gama Camera including accessories. | 1 |

Consignee: Tertiary Cancer Care Centre at Government Medical College, Kozhikode.

Part II: Required Delivery Schedule:

a) For Indigenous goods or for imported goods if supplied from India or for the imported goods against site draft:

120 days from date of Notification of Award to delivery at consignee site. The date of delivery will be the date of delivery at consignee site (Tenderers may quote earliest delivery period).

b) For Imported goods directly from foreign:

120 days from the date of Notification of Award. However, the delivery period in LC will be given for 90 days from the opening of LC. The date of delivery will be the date of receipt of goods by the consignee. (Tenderers may quote the earliest delivery period). The LD charges as applicable will be deducted for the delay after 120 days from NoA date.

Special Note for imported goods:

1. The required documents for LC opening and other related documents (like NoA Acceptance, Contract Agreement signing, PBG submission, Performa Invoice, any other related documents etc.) satisfied/ acceptable to the purchaser, fulfilling Government/ bank requirement shall be submitted by the supplier to HLL within 30 days from the NoA date. It is the responsibility of the supplier to ensure that the consignment is delivering within 120 days from the NoA date.

Part III: Scope of Incidental Services:

Installation & Commissioning, Supervision, Demonstration, Trial run and Training etc. as specified in GCC Clause 13

Part IV:

Turnkey (if any) as per details in Technical Specification.

Part V:

Comprehensive Maintenance Contract (CMC) as per details in Technical Specification.

Part VI:

Required Terms of Delivery and Destination.

a) For Indigenous goods or for imported goods if supplied from India or for imported goods against site draft:

At Consignee Site – Specified in the List of Requirements

Insurance (local transportation and storage) would be borne by the Supplier from ware house to the consignee site for a period including 3 months beyond date of delivery

b) For Imported goods directly from abroad:

The foreign tenderers are required to quote their rates on DDP basis at consignee site.

The shipping arrangements shall be made in accordance with the instruction of Ministry of Shipping & Transport, New Delhi, India as detailed in Annexure 1 at Section XIX.

Insurance (local transportation and storage) would be extended and borne by the Supplier from ware house to the consignee site for a period including 3 months beyond date of delivery.

Destination/Consignee details are given in Section XXI

Section – VII

Technical Specifications

FOR

SPECT/CT SYSTEM

PART A: GENERAL

- GENERAL:

- A latest technology Dual Headed variable angle SPECT (CT) Gamma Camera system capable of performing static imaging along with all other planar, SPECT, WB SPECT, Cardiac SPECT and whole body imaging applications with fast rotation CT for attenuation correction and functional anatomical image fusion  .

- System should be capable to perform CT based attenuation correction for nuclear medicine images.

- System should be supply along with Image fusion software & hardware.

- Manufacturer should provide a comprehensive list of users of dual head SPECT (CT) gamma camera in India and their performance profile.

- Number of installation and company profile:  Company should preferably be manufacturer of the equipment and in case non-manufacturing vendors (third party supplier) ,the same should be mentioned clearly in bid document.  In case of third party supplier, the liabilities of principal manufacturing company and third party supplier need to be clearly demarcated and ensured.

- Bidder should carryout installation and commissioning on turnkey project basis.

- The product should be US FDA certified. Vendor should provide type approval certificate of quoted model from AERB and copy should be attached with technical bid.

PART B: GAMMA CAMERA

1. GANTRY:

- Gantry design should be wide open to avoid claustrophobic  imaging  acquisition capability with clockwise and anticlockwise movement.

- Height, width and depth should be adequate to conveniently locate the gantry in existing space available in the department.  In order to access this, bidders are requested to visit the site before submitting the tender.

- Auto contouring: The detectors should be equipped with automatic body countering (ABC).

- Gantry should have emergency stop buttons.

- Persistence scope mounted on the gantry or wall for continuous display of patient position and gantry parameters (with LCD Color Display)

2. DETECTORS:

- Large field of view with UFOV of 45 X 35 cm or more to enable adequate patient breadth coverage and diagonal field of view of 50 cm or more.

- Crystal thickness should be 9.5 mm (3/8”).

- Number of PMTs should be 55 or more per detector with 1 ADC per PMT (True digital detector).

- System should have facility for automatic correction for energy, linearity and uniformity.

- As per NEMA NU-1, 2007 standard please specify values of the offered system.

  - Intrinsic Spatial Resolution:  minimum requirement 3.4 mm or better than that.

  - Intrinsic (integral) Flood Field Uniformity: minimum requirement: 2.8 % or better than that.

  - System Spatial Resolution (LEHR at 10 cm without scatter): minimum requirement 7.5 mm or better than that.

- Collimator Specification:

- Low Energy High Resolution (LEHR) – One pair

- Hi Energy General Purpose (HEGP) - 0ne pair

- Medium energy general purpose (MEGP) one pair.

Collimator exchange cart for each set of collimators has to be supplied.

- PATIENT BED:

- Single universal table for all studies i.e., Planar, SPECT and Whole body imaging and CT images.

- Table top should be composed of low attenuation carbon fiber covered by mattress pad and  straps and should be able to withstand at least 180 Kg of body weight

- Whole body imaging covering should not be less than 190 cm or more.

- Table should have facility for lowering the height to facilitate easy patient transfer and should be movable to permit imaging for sitting, standing, stretcher/wheel chair patients

Adjustable head positioning pallets, injection arm rest, Cardiac arm rest, leg support, patient restraint & support

- System compatible ECG Trigger to be provided.

- ACQUISITION STATION

- One Acquisition station independent of main processing unit capable of data acquisition in Cardiac SPECT, Gated SPECT, Static, dynamic, multi-gated, whole body scanning, SPECT and Gated SPECT.

- High performance latest PC with multi tasking operating system having minimum of 4 GB RAM, 2.8 GHz processor speed, 1 TB or more SCSI hard drive and high resolution flat panel LCD monitor of minimum of 19” size. It should also have CD and DVD combo drive preferably with writer facility

- Image acquisition and data display should be from 64 x 64 matrix up to 256 x 1024 matrix.

- Acquisition termination by preset time, preset count or manual stop.

- Pre-defined acquisition protocols as well facility for user to configure his own customized protocols. Should provide ECG Gating Device with R wave Display

- DATA ACQUISITION

- Cardiac SPECT and  First pass studies.

- Cardiac shunt studies.

- Gated Cardiac Studies (Automatic/Manual).

- Ventilation/Perfusion studies.

- Clinical Cerebral Studies including HMPAO-SPECT.

- Renal Q/F Analysis and MAG-3 protocol DTPA protocol.

- Liver/Abdominal Studies.

- Whole Body Studies – anterior and posterior, dual intensity.

- Whole Body SPECT facility.

- Thyroid studies.

- Quantitative Myocardial SPECT package.

- Gallium Whole Body anterior/posterior, SPECT acquisition, reconstruction and Analysis.

- Motion correction.

- Gall Bladder Ejection fraction.

- Comprehensive DICOM access shall be available. MMPS shall be available.

- MULTI  MODALITY PROCESSING WORK STATION

- High performance latest Multi tasking PC with full DICOM readiness for image

transfer print etc.

- Minimum of 8 GB RAM, 3.0 GHz processor speed and 500 GB SCSI hard drive

- High resolution flat panel LCD monitor of minimum of 21” size.

Standard DVD & CD combo drive for data Archiving and Retrieving with writer facility for both.

- CLINICAL APPLICATIONS SOFTWARE:

- Specific software should be provided to apply CT based attenuation correction on cardiac study and other nuclear medicine study.

- Processing workstation should have following quantification program

  - Thyroid uptake calculation,

  - GFR and renal function calculation, VUR graph and calculation

  - Lung segmental analysis right and left ratio.

  - Gastric emptying, reflux  calculation

  - Salivary gland emptying calculation software

  - Image fusion, adding, subtraction software

  - Graph creating, manipulation software

- Packages should including Display Analysis software, 3-D volume rendering display with maximum intensity projection (MIP), cine review capability, Image manipulation tools and curve generation and manipulation tools..

- Image output format should include JPEG, TIFF, AVI and in DICOM format.

- NEMA NU 1 – 2007 or latest display software QC test program.

- Anatomical and functional data co-registration software for exact image localization (SPECT/CT fusion software).

- Apart from inbuilt SPECT-CT fusion software, separate software, for fusion of imported CT and MRI data with SPECT data, has also to be provided.

- Transmission attenuation correction software to generate attenuation map complete access to the existing compiler and editor in which the software is provided in order to modify any part of the existing program as and when needed.

- SPECIAL CARDIAC APPLICATIONS:

- Complete Cardiac package including First Pass EF and Cardiac Shunt quantification studies,gated equilibrium, MUGA SPECT, Myocardial Perfusion (Planner and SPECT) and gated SPECT tomography including Emory Cardiac Tool-box (ECT)or equivalent validated software, QGS, QPS and QBS should also be provided.

- Addition to the Filter back projection algorithm also require 3D-OSEM reconstruction algorithm to enhanced image contrast and enhanced signal to noise ratio.

- Latest Special software or hardware available with the respective vendor should offer as a standard features which helps to enhance the cardiac SPECT images quality and to complete the study in minimum time. Provide specific detail on the offered cardiac package.

- QUALITY CONTROL TOOLS:

- CT quality assurance phantom for contrast resolution, radiation safety and image uniformity and pixel noise etc.

- QC software (CT) for verifying alignment of the table position between SPECT and CT acquisition

- SPECT phantoms, point sources, CT phantom and KV/mA meter to permit comprehensive quality assurance program both for SPECT and CT. Four Quadrant BAR phantom for rectangular detector size compatible with the detector size of the camera.

- Cobalt 57 flood source of at least 15 mCi strength (on the day of the delivery to the institute) for rectangular field of the size adequate for the gamma camera to be supplied.

- Intrinsic and  system resolution phantom.

- System count rate performance phantom.

- QC software package (NEMA NU2-2007 or latest protocol) with documentation.

PART C: COMPUTED TOMOGRPHY SYSTEM

- GENERAL:

- SPECT system should have in-build multi slice CT scanning facilities to generate attenuation correction map to apply on Nuclear Medicine images for body attenuation correction and to get accurate localization by fusing the SPECT and CT images.

- CT system should provide minimum radiation dose to the patient to generate attenuation correction map and to get high spatial resolution CT images.

- Specify CTD body dose for body attenuation and body localization.

- X-RAY TUBE & GENERATOR:

Specify following for your system

- Generator Power rating (kW):

- KV range: Specify & enumerate KV & mA :

- Specify all parameters related with tube assembly

- Anode heat storage capacity

- Mention the details of the cooling systems

- specify the focal spot Size and number,Automatic selection of spot size

- Tube current (mA)  and Tube Voltage (kV)

- Rotation time: Second

- Temporal resolution with heart view CT option: micro Second

- Single continuous spiral scan time: Seconds

- Gantry:

Aperture should be 70cms or more to facilitate entire patient population.

Specify the entire range of rotation times for full 360 degrees.

It should be fast rotation CT scanner.

Minimum scan field of view (FOV) should be 50cms or more.

Controls should be located on both sides within easy reach of the operator.

4)SPIRAL CT:

Scan time for complete 360 degree scan should be 1 sec or less.

Minimum slice thickness should be less than 3mm to match with SPECT slices.  Gapless spiral length 100cms or more.

Specify range and selectable slice thickness available.

Pitch factor (volume pitch) should be variable between 0.75 to 1.5 or more and should be user selectable. Give details of all pitch selections.

Rotation times:  specify the minimum and maximum with the range of pitch. Combinations of pitch, slice thickness and spiral length should be easily selectable by the user.

The system should optimize the radiation dose and resolution for each selection.

- TOPOGRAM:

Specify the range of length.

Scan time: Give details

Views:  should be possible in frontal and lateral views.

5.DATA ACQUISITION SYSTEM:

- Detector: Please specify the detector design and type of detector used.

- Give details of the detector electronics, number of rows with their thickness, number of

elements, channels per slice, number of projections, maximum number of discrete slices/rotation.

- System display of radiation dose per examination

- There should be in-built protocols for all pediatric applications in order to reduce the

effective dose.

6.IMAGE RECONSTRUCTION:

- Specify the reconstructed slice width, scan field and reconstruction field.

- Specify the scan time to perform breath hold and tidal breathing CT scan to cover 14 cm (heart region)

- IMAGE DISPLAY MONITORS (common with the Gamma Camera):

Monitors  should be flat screen to last full life of the scanner with continuous and uninterrupted use without significant deterioration of display quality.

Minimum numbers of monitors to be provided are 3. One for the examination room, one for the control room and one for post processing console/ work station

Size of monitors should be more than 18":  Full flat screen LCD display with resolution of 1024 x 1024 or more

Pixel size should be minimum 0.3 mm or less.

Afterglow persistence is undesirable. Please give details.

7.CONSOLE AND PROCESSING WORKSTATION:

  - A common console for Gamma camera and CT system

  - Processing workstation should have standard CT viewer to display the CT images with batch viewing and oblique slices (slab view)

  - Integrated Operator console should perform all functions such as CD writing, acquisition protocol customization, scheduling, protocol selection, all standard evaluation applications.

  - Acquisition system should accept DICOM patient work list.

  - Please specify all post processing software which will run on the workstation and not on the operating console.

  - Independent Advanced Workstation/ Console X 2 Nos.

  - Capable of simultaneous acquisition and post processing, preferably swappable with the main console should be provided. These should have hard disk storage of 1 TB or more with options for provided enhancement of storage capacity and future upgrades. Give details of the hardware of workstations/ consoles to be provided.

The following software is to be provided:

Real time reconstruction and immediate display of images in 1024 x 1024 matrix parallel to acquisition in spiral mode to be provided.

8.IMAGE EVALUATION TOOLS:

Parallel evaluation of multiple regions of interest in circle, irregular and polygonal form.

Complete statistical evaluation software for axial & 3D images. Blood flow quantification packages for perfusion  studies, Give details.

Profile cuts, horizontal, vertical, oblique.

Annotation, labeling, online measurements, ROI etc.

Dynamic evaluation of contrast enhancement in organs and tissue, calculation of time density curves, peak enhancement images, time to peak images, image subtraction etc.

2D image zoom, pan, manipulate, averaging reversal of grey scale values, mirroring, image filter functions, advanced imaging algorithms.

Fusion of morphological data obtained on CT & MR.

All above functions should be instantaneous and interactive.

On line patient registration facility, transfer of information from HIS/RIS via DICOM and emergency registration.

9.CT IMAGE QUALITY:

Following parameters must be specified clearly as these are considered crucial for system comparison.

- Specify visual and statical low contrast resolution.

- Specify surface dose, mAs, slice thickness and HU used to determine the

low contrast resolution.

- Spatial resolution should be more that or equivalent to 15lp/cm @ 10

- Trans axial field of view should be 45 cm or more.

- Slice thickness should be 0.5 mm or less than that.

10.MAGE ARCHIVING

1.IMAGE DOCUMENTATION AND ARCHIVING: A dry laser imager with digital interface and control integrated with the main console. Camera should print on 8"X10", 11"X14", 14"X17" with flexible formatting on 14" x 17" film.

- The dry laser imager should have software controlled operational capabilities with a resolution of at least 600 DPI or more.

- DICOM connectivity should be optimized for networking (send and receive facility) with other imaging systems. It should have sufficient memory to store images from other systems (CT, MR, DSA and U.S.) connected to it.

- Filming should be parallel to other activities including independent scanning, documentation and post processing. Should be possible to configure image text.

- Background filming should be possible.

- Full specifications of the laser camera to be offered must be provided.

- DOCUMENTATION OF COLOUR IMAGES: A high resolution laser colour printer (min 1200 dpi) to be provided for printing of colour angio, perfusion, 3D images etc. Print media or suitable glazed paper and cartridge for the printer to be provided for approximately 500 print outs per year during the warranty period

.

- ARCHIVING:

    - DVD DICOM recorder with at least 4 GB storage per side per disc with facility for quick retrieval of images should be provided 25 DVDs per year for the entire warranty period should be quoted.

    - The archiving system should have a patient data base to track archived images. Include PC with latest configuration.

- IMAGE TRANSFER/NETWORK:

- System should be fully DICOM 3 compatible for DICOM send/receive/query/ retrieve, basic print connectivity to any network, and minimum 5 PC’s. It should be possible to integrate the system to any PACs, HIS & RIS for transfer of CT images to the Departments/sections (like OPD chambers, wards, Board Room, OT, RT section, doctors rooms etc) Medical College Hospital and should be carried out by the vendor at their cost.

- Vendor should enable live image display in the conference area of the depth, directly from the system using slave monitors.

- Global modem to allow on-site connectivity and maintenance.

- All necessary interconnecting interfaces, cables, modules and other hardware and software to fully integrate the system and seamlessly merge the new facility with the current and future workflow of the Institute.

###### PART  D :  WARRANTY & AFTER SALES SERVICE

The vendor shall give a comprehensive, on-site warranty from the Principals for the first three years from the date of commissioning of the entire system (inclusive of vacuum and non-vacuum parts and all the locally supplied items including consumables like batteries of the UPS, printer cartridges etc). Rates for comprehensive annual maintenance for the next five years’ period shall be quoted separately on annual basis . Replacement cost of The Digital detectors, crystals, PM Tubes and the X-ray tube (for the Gamma Camera and for the CT) CAMC period should be covered in CAMC price. Pro-rata warranty is not acceptable. Maintenance Contract offered by the local agents is also not acceptable.

A written agreement valid for five years for complete unit, signed by the manufacturer, shall be provided to guarantee that failures in materials and workmanship that occur within the warranty period will be corrected.

Warranty, guarantee and service are considered as part of the bid specification. The contractor shall provide complete and specific details of maintenance operations performed under service contracts

The service engineer, representing the vendor shall be thoroughly familiar and experienced with all aspects of SPECT CT system.

The manufacturer shall give an unqualified commitment that the maintenance and guarantee of the equipments shall in no way be affected by any change in its administrative or trade set up like the change in the authorized agent or any merger, amalgamation or separation of the company or any of its constituents etc for that matter.

During the period of maintenance warranty, the supplier shall give a commitment for 98% uptime of the equipment as well as the accessories, calculated on annual basis, with penalty equivalent to double the amount of daily cost of the unit for each day’s delay in proper functioning of the unit beyond 5% downtime per annum (calculated on the basis of the total cost of the AT divided by the total life span of the equipment taken as equivalent to 3000 working days). That means a maximum total 15 (fifteen) days of downtime shall be permitted in a span of every 12 months, without invoking the penalty clause. However, if the machine lies non-functional for a period of more than one week continuously at any stage, the same penalty will be imposed even if 98% uptime clause is being met with for the given calendar year. Any bid without agreeing to this warranty and penalty clauses will be summarily rejected.

Any recall of hardware or software of the equipment by manufacturer or regulatory body like US FDA, shall have to be replaced at this institution by the vendor immediately, free of costs, whatsoever.

Detailed Technical Catalogues (Service Manuals & User’s Manuals-original sets) shall be provided for all the equipment including accessories.

All vendors shall quote their bids showing the break up prices instead of a composite price, both for the equipments as well as for the turnkey work. The break up prices of the equipments, and their major components/accessories/ spares shall be clearly and unambiguously stated to facilitate comparison of prices on a common platform.

PART  E  : TRAINING AND AFTER SALE SERVICE

- Training as per requirement.  In order to fully and optimally utilize the equipment, training for two personnel Technical&Clinical) at least for a period of 2 weeks shall be provided in the factory as well as in a reputed international/domestic institution by the vendor.  Practical training to at least two persons shall be provided at a functioning site within India.  All the financial commitments for the trainings shall be met by the vendor. Besides, the supplier shall also provide complete on-site training to all the staff of the department.

- On site application training in phases to technical and clinical staff initially and also on as and when required basis.

- A fully trained local resident engineer to be posted to ensure optimal uptime and effective and prompt maintenance  services subsequent to satisfactory installation.

- The acceptance of the installation shall be subject to satisfactory handing over of the system to the department and certificate to this effect to be issued by the institute.

- All the operating and service manual in duplicates to be provided by the vendor at the time of handing over the machine.

- Books and journals:  Four books and two journals (subscribed for 5 years) shall be provided free of cost as per the choice and requirement of the institute.

PART  F :  ESSENTIAL ACCESSORIES TO BE SUPPLIED

| No | Items | Units |
| --- | --- | --- |
| 1 | Dry Laser Imager | 1 |
| 2 | Lead Aprons | 2 |
| 3 | L Bench NM- | 2 |
| 4 | Stainless Steel Syringe carrier having lead lining of 4mm | 4 |
| 5 | Contamination monitor | 1 |
| 6 | Gun Monitor (ion chamber) | 1 |
| 7 | Survey Meter GM Type | 1 |
| 8 | Lead Bricks set of 54 Bricks | 2 |
| 9 | Lead lined Waste Box - NM | 2 |
| 10 | Lead lined decay Drum - NM | 2 |
| 11 | Hot lab tongs and forceps | 4 |
| 12 | Colour laser Printer | 1 |
| 13 | Online UPS with maintenance free batteries for whole system with 30 min back up. | 1 |
| 14 | PC based Cardiac treadmill from reputed manufacturer. | 1 |
| 15 | Infusion pump with 5 to 50ml range syringe capacity | 1 |
| 16 | Radioisotope calibrator for beta and gamma activity measurement | 2 |
| 17 | Pocket dosimeters | 2 |
| 18 | Syringe holders | 2 |
| 19 | Gamma zone monitor | 2 |
| 20 | Auto dispenser for radio pharmaceuticals | 1 |
| 21 | Cutie Pie for Gamma detection, | 1 |
| 22 | Additional flat table top with complete hardware and software for radiotherapy planning. | 1 |
| 23 | Vial opener (Open source) | 2 |
| 24 | Aerosol generator (lung ventilation study) | 1 |
| 25 | Hand foot monitor | 1 |
| 26 | Sources for calibration | 1 |
| 27 | Lead lined Fume Hood with exhaust system, Airfoil, Air Baffle and adjustable slots, sliding sash, Bypass slots, exhaust duct and damper sinks | 2 |
| 28 | Vacuum cleaner | 1 |
| 29 | Scintimammo pads | 1 |

Note: Price of the each accessory should be mentioned separately in the price bid.

PART  G  :   UPGRADING REQUIREMENTS

Upgrades:  A free  Comprehensive software  upgrade guarantee for three years from commissioning of the scanner must be provided for free of cost.

PART H: TURNKEY WORKS

Bidder should carryout installation and commissioning of SPECT Gamma Camera on turnkey project basis in the existing nuclear medicine department. The Turnkey works including minor civil works, inside partitions, interior works, flooring, false ceiling, painting, electrification works, supply of furniture, electrical fittings, air-conditioning etc. The room layout plan is attached. However bidders are strongly advised to verify the site conditions and dimensions before submitting the bid.Bidder shall carry out  all necessary work required for satisfactory working of the equipment and mandatory requirement as per AERB guidelines.

1. Civil work:

All the civil work has to be done as per the final plan approved by the purchaser. Bidder should carryout civil works fully complying with directorate of Radiation Safety guidelines, Govt. of      Kerala wherever it is applicable

Flooring-

- Seamless vinyl flooring

- Approved Make: Armstrong

Painting

- Plastic Emulsion Paint two coats over 2 coats of wall putty

- Paints should be of standard make like Asian paints or Dulux

False Ceiling

- Acoustical tile for ceiling, with light weight insulating material of high quality, supported on grid or finished seamless with support above ceiling. Finished with white paint or powder coated with white paint, if metallic. Ceiling height to suit the equipment mount and clearances.

2. ELECTRICAL WORKS

The supply line will be provided by the Institute up to the distribution panel with existing cables.  The distribution panel should have switch gear of makes like Siemens/ L & T and shall be provided by the vendor. Bidder has to carryout electrical works as per PWD guidelines. Dedicated electrical grounding has to be provided by the vendor.

The electrical work shall include the following:

- Wiring –All interior electrical wiring- with main distribution panel board, necessary MCBs, DB, joint box, switch box etc. the wires shall be of copper of different capacity as per the load and should be renowned make like Finolex, Polycab.

- Switches light and power points should be of modular type and of standard make

- General lights –Mirror optical type 1X28 W or 2X28 W/CFL fittings 2X36, 3X36 W with electronic ballasts  Philips/Crompton/Kesselec-Schreder/Wipro make.

- MCBs/ACBs/MCCBs should be of standard make like SIEMENS/ABB/ L & T.

- Telephone cables should be of standard make like FINOLEX  cables

3. PLUMBING WORKS

All water pipes and fittings shall be of high density polythene of approved and standard make.  The gratings shall be brass chrome plated.All plumbing accessories should be of standard make like Jaguar.

4. AIR CONDITIONING

Ductable A/C Package /Split AC units of 5 star rating may be used according to room requirement and suitability. Humidity control should be effective to eliminate moisture condensation on equipment surface.

Environment specifications:

a) Humidity range:  Relative humidity 40% and 60%  in all areas except equipment room which shall be as per requirement of the equipment.

b) Temperature ranges:  22 + 2° C in all areas except equipment room which shall be as per requirement of the equipment.

c) Air conditioning load:The heat load calculations and maintaining the desired temperature and humidity shall be the responsibility of the bidder

5. FURNITURES

All furniture items should be of standard make like Hermen Miller, Godrej etc.

a) Tables to keep workstations and processing stations.

b) Revolving chairs height adjustable ,medium-back with hand-rest in the Control room, and viewing area.

c) Name boards for all rooms

d) Suitable lockers and shelves to keep accessories

f) Dustbins (plastic with lid) to be provided as required.

6.FIRE FIGHTING

Bidder should provide effective fire fighting system. Bidder should provide two Dry CO2 cyliners-2 kg with essential accessories. Cylinders should be certified by respective regulatory board

| SN | ITEM | Approved Make |
| --- | --- | --- |
| A | CIVIL WORKS |  |
| 1 | Vinyl Flooring | Armstrong |
| 2 | Paint | Dulux/Asian |
| B | PLUMBING ITEMS | Kohler, Jaguar |
| C | SANITARY ITEMS | Cera,Hindware,Parryware |
| D | ELECTRICAL WORKS |  |
| 1 | Cables | Fonlex,Havells |
| 2 | Switches | Legrand, L&T, Crabtree |
| 3 | Distribution Box, Breakers | Legrand, L&T, Siemens |
| 4 | Lighting | Philips, GE |
| E | AIR CONDINTIONING | Bluestar, Voltas, Daikin |
| F | FURNITURE | Hillrom, Hermen Miller,Godrej |

Bidder should mention break up price for all turnkey separately in the price bid.

| SN | TYPE OF WORK | Total cost in INR |
| --- | --- | --- |
| 1 | CIVIL WORKS |  |
| 2 | AIR CONDITIONING |  |
| 3 | ELECTRICAL WORKS |  |
| 4 | FIRE FIGHTING | ` |
| 5 | PLUMBING |  |
| 6 | FURNITURE |  |

Annexure-1: Site Drawing

SECTION-VII

TECHNICAL SPECIFICATIONS

GENERAL TECHNICAL SPECIFICATIONS FOR EQUIPMENT

GENERAL POINTS:

1.    Warranty:

- Three years Comprehensive Warranty as per Conditions of Contract of the TE document for complete equipment and Turnkey Work from the date of satisfactory installation, commissioning , trial run & handing over of equipment to Institution/Consignee/ User.

- 98% up time Warranty of complete equipment with extension of Warranty period by double the downtime period on 24 (hrs) X 7 (days) X 365 (days) basis.

- All software updates and changes should be provided free of cost during Warranty & CMC period.

- After Sales Service:

After sales service centre should be available at the city of Hospital/Institution/Medical College/Organisation/Consignee on 24 (hrs) X 7 (days) X 365 (days) basis. Complaints should be attended properly, maximum within 24 hrs. The service should be provided by Tenderer/Indian Agent.  Undertaking by the Principals that the spares for the equipment shall be available for at least 10 years from the date of supply.  If the manufacturer /agent do not have the service centres in India they will have to set up the same within 45 days after award of contract.

- Training:

On Site training/ at Principals site to Technicians/ staff is to be provided by Principal/ Indian Agents (if they have the requisite know-how) for operation and maintenance of the equipment to the satisfaction of the consignee.

4.     Annual Comprehensive Maintenance Contract (CMC) of subject equipment with Turnkey:

a)	The cost of Comprehensive Maintenance Contract (CMC) which includes preventive maintenance including testing & calibration as per technical/ service /operational manual of the manufacturer, labour and spares, after satisfactory completion of Warranty period may be quoted for next 5 years on yearly basis for complete equipment (including other vacuumatic parts wherever applicable) and Turnkey (if any). The supplier shall visit each consignee site as recommended in the manufacturer’s technical/ service /operational manual, at least once in three months during the CMC period

b)	The cost of CMC may be quoted along with taxes applicable on the date of Tender Opening. The taxes to be paid extra, to be specifically stated. In the absence of any such stipulation the price will be taken inclusive of such taxes and no claim for the same will be entertained later.

c)	Cost of CMC will be added for Ranking/Evaluation purpose.

d)	The payment of CMC will be made on six monthly basis after satisfactory completion of said period, duly certified by end user on receipt of bank guarantee for 2.5 % of the cost of the equipment as per Section XV valid till 2 months after expiry of entire CMC period.

e)	There will be 98% uptime warranty during CMC period on 24 (hrs) X 7 (days) X 365 (days) basis, with penalty, to extend CMC period by double the downtime period.

f)	During CMC period, the supplier is required to visit at each consignee’s site at least once in 6 months commencing from the date of the successful completion of warranty period for preventive maintenance of the goods.

g)   All software updates and changes should be provided free of cost during CMC.

h)	Failure of the above [4. e) to 4. g)] by the supplier, may lead to the forfeiture of the Bank Guarantee for Annual CMC.

i)	The payment of CMC will be made as stipulated in GCC Clause 21.

##### Turnkey:

Turnkey is indicated in the technical specification, ifrequired. The Tenderer shall examine the existing site where the equipment is to be installed, in consultation with the consignee.  The Tenderers to quote prices indicating break-up of prices of the Machine and Turnkey Job of each equipment. The Turnkey costs may be quoted in Indian Rupee will be added for Ranking Purpose.

The taxes to be paid extra, to be specifically stated. In the absence of any such stipulation the price will be taken inclusive of such duties and taxes and no claim for the same will be entertained later.

The Turnkey Work should completely comply with AERB requirement, if any.

Section – VIII

Quality Control Requirements

(Proforma for equipment and quality control employed by the manufacturer(s)

Tender Reference No.

Date of opening

Time

Name and address of the Tenderer:

Note:  All the following details shall relate to the manufacturer(s) for the goods quoted for.

01	Name of the manufacturer

a.	full postal address

b.	full address of the premises

c.	telegraphic address

d.	telex number

e.	telephone number

f.	fax number

02	Plant and machinery details

03	Manufacturing process details

04	Monthly (single shift) production   capacity of goods quoted for

a.	normal

b.	maximum

05	Total annual turn-over (value in Rupees)

06	Quality control arrangement details

a.	for incoming materials and bought-out components

b.	for process control

c.	for final product evaluation

07	Test certificate held

a	. type test

b	. BIS/ISO certification

c	. any other

08	Details of staff

a.	technical

b	skilled

c	unskilled

Signature and seal of the Tenderer

Section – IX

Qualification Criteria

  - The Tenderer must be a Manufacturer or its authorized Agent.

02.	(a)	The Manufacturer should have supplied and installed in last Five years from the date of Tender Opening, atleast 100% of the quoted quantity of the similar equipment meeting major specification parameters which is functoning satisfactorily. The foreign Manufacturer satisfying the above criteria should also have supplied and installed in last Five years from the date of Tender Opening, 100% of quoted quantity of similar model which is functoning satisfactorily any where outside the country of manufacture.

02.	(b)	The Tenderers quoting as authorized representative of the manufacturer meeting the above criteria 02 (a) should have supplied and installed in last Five years from the date of Tender Opening, atleast 100% of the quoted quantity of similar equipments which is functoning satisfactorily, any where in India of the same manufacturer.

Note

- In support of 2 (a) & 2 (b), the Tenderer shall furnish Performance statement in the enclosed Proforma ‘A’.

The manufacturer as well as the Tenderer/ Indian Agent shall furnish Satisfactory Performance cum installation Certificate in respect of above, duly translated in English and duly  notarized  in the country of origin, alongwith the tender.

- The Tenderer shall furnish a brief write-up, packed with adequate data explaining and establishing his available capacity/capability (both technical and financial) to perform the Contract (if awarded) within the stipulated time period, after meeting all its current/present commitments. The Tenderer shall also furnish details of Equipment and Quality Control in the enclosed Section VIII.

- Notwithstanding anything stated above, the Purchaser reserves the right to assess the Tenderer’s capability and capacity to perform the contract  satisfactorily before deciding on award of Contract, should circumstances warrant such an assessment in the overall interest of the Purchaser.

- The Purchaser reserves the right to ask for a free demonstration of the quoted equipment at a pre determined place acceptable to the purchaser for technical acceptability as per the tender specifications, before the opening of the Price Tender.

###### PROFORMA ‘A’

# PROFORMA FOR PERFORMANCE STATEMENT

(For the period of last five years from the date of tender opening)

Tender Reference No.	: _________________________________

Date of opening	: _________________________________

Time	: _________________________________

# Name and address of the Tenderer	: _________________________________

# Name and address of the manufacturer	: _________________________________

| Order placed by (full address of Purchaser/Consignee) | Order number and date | Description and quantity of ordered goods and services | Value of order (Rs.) | Date of completion of Contract |  | Remarks indicating reasons for delay if any | Have the goods been functioning Satisfactorily (attach end user certificates as per format annexed)** |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | As per contract | Actual |  |  |
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |

Signature and seal of the Tenderer

** The documentary proof will be a certificate from the consignee/end user with cross-reference of order no. and date in the certificate along with a notarized certification authenticating the correctness of the information furnished. If at any time, information furnished is proved to be false or incorrect, the earnest money furnished will be forfeited

FORMAT OF PERFORMANCE CERTIFICATE

To whom it may concern

Date__________________

Certified that M/s____________________________________(name & address of manufacturer) supplied us ________________________Nos(indicate quantity) of equipment,____________________________________   (indicate name of the equipment) against our order no_________________________dt___________________________(please indicate order no & date as figuring in the performance statement).The equipment was installed,commissioned and handed over to us __________________(indicate date) & since then the equipment is has been working to our entire satisfaction.

Place:__________________                                                           Name & Designation of the Date:__________________                                                            officer with seal__________

_______________________

(in capital letters)

### Section – X

TENDER FORM

Date__________

To

The DVP (Tech.), HLL Lifecare Ltd, (A Government of India Enterprise), Infrastructure Development Division, “Adarsh” TC 6/1781, Vettamukku, Thirumala P.O, Thiruvananthapuram-695006

Ref. Your TE document No. _________dated ___________

We, the undersigned have examined the above mentioned TE document, including amendment/corrigendum No. __________, dated ________ (if any), the receipt of which is hereby confirmed. We now offer to supply and deliver___________ (Description of goods and services) in conformity with your above referred document attached herewith and made part of this tender.

If our tender is accepted, we undertake to supply the goods and perform the services as mentioned above, in accordance with the delivery schedule specified in the List of Requirements.

We further confirm that, if our tender is accepted, we shall provide you with a performance security of required amount in an acceptable form in terms of GCC clause 5, read with modification, if any, in Section - V – “Special Conditions of Contract”, for due performance of the contract.

We agree to keep our tender valid for acceptance as required in the GIT clause 20, read with modification, if any in Section - III – “Special Instructions to Tenderers” or for subsequently extended period, if any, agreed to by us. We also accordingly confirm to abide by this tender up to the aforesaid period and this tender may be accepted any time before the expiry of the aforesaid period. We further confirm that, until a formal contract is executed, this tender read with your written acceptance thereof within the aforesaid period shall constitute a binding contract between us.

We further understand that you are not bound to accept the lowest or any tender you may receive against your above-referred tender enquiry.

We confirm that we do not stand deregistered/banned/blacklisted by any statutory Authorities as per govt. rules/procedures..

We confirm that we fully agree to the terms and conditions specified in above mentioned TE document, including amendment/ corrigendum if any

(Signature with date)

(Name and designation) Duly authorised to sign tender for and on behalf of

SECTION – XI PRICE SCHEDULE

A)	PRICE SCHEDULE FOR DOMESTIC GOODS OR GOODS OF FOREIGN ORIGIN LOCATED WITHIN INDIA

| 1 | 2 | 3 | 4 | 5 |  |  |  |  |  |  | 6 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Schedule | Brief Description of Goods | Country of Origin | Quantity (Nos.) | Price per unit (Rs.) |  |  |  |  |  |  | Total Price (at Consignee Site) basis (Rs.) 4 x 5(g) |
|  |  |  |  | Ex - factory/ Ex -warehouse /Ex-showroom /Off - the shelf (a) | Excise Duty (if any) [%age & value] (b) | Sales Tax/ VAT(if any) [%age & value] © | Transportation Charges (d) | Insurance charges for a period including 3 months beyond date of delivery, loading/ unloading and Incidental costs till consignee’s site (e) | Incidental Services (including Installation & Commissioning, Supervision, Demonstration and Training) at the Consignee’s site (f) | Unit Price (at Consignee Site) basis (g) =a+b+c+d+e+f |  |

Total Tender price in Rupees: ___________________________________________________________________________________________________

In words: ____________________________________________________________________________________________________________________________

Note: -

- If there is a discrepancy between the unit price and total price THE UNIT PRICE shall prevail.

- The charges for Annual CMC after warranty shall be quoted separately as per Section – XI – Price Schedule C

Name________________________

Business Address________________________

Place: ___________________________ 	Signature of Tenderer________________________

Date: _________________________ 	Seal of the Tenderer________________________

SECTION – XI PRICE SCHEDULE

B)	PRICE SCHEDULE FOR GOODS TO BE IMPORTED FROM ABROAD

| 1 | 2 | 3 | 4 | 5 |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Schedule | Brief Description of Goods | Country of Origin | Quantity (Nos.) | Price per unit (Currency) |  |  |  |  |  |  |  |  |  |  |
|  |  |  |  | Gross FOB price at sea/air port of Lading (inclusive of Agency Commission) | Amount and percentage of Agency Commission ** | Net FOB (excluding Agency Commission) (a-b) | Insurance & Freight | Net CIP by Air/ Sea at the port of entry (c+d) | Custom Duty amount as % of Net CIP (amount with CDEC as applicable) ** | Custom Clearance & Handling Charges ** | Loading/ unloading, inland transportation, insurance as per Clause 11 of GCC & incidental cost till consignee’s site ** | Installation commissioning,supervision. Demonstration & training at the consignee’s site ** | Unit price on DDP basis at consignee’s site |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  | (i) In Indian Rupees (b+f+g+h+i) | (ii) In foreign currency (e) |
|  |  |  |  | (a) | (b) | (c) | (d) | (e) | (f) | (g) | (h) | (i) | (j) |  |

Total price at Consignee’s site

** (i)       In Indian Rupees            column 4 x (b+f+g+h+i) Rs_______________________________________________(In figures and words)

(ii)      In foreign currency         column (4 x e)_________________________________________________________ (In figures and words)

________________________ _______________________________________________________________________________

Note: -

1. The Tenderer will be fully responsible for the safe arrival of the goods at the consignee site in good condition as per terms of DDP and INCOTERMS

2. The bidders break up of prices under various columns are for comparison of prices up to delivery of goods at consignee’s site for tender evaluation.

3. The quoted price should be supported with original proforma invoice from the foreign manufacturers. The proforma invoice should indicate the percentage of agency commission included in the FOB

prices. Indian Agent to be paid in Indian Currency.

4. All the components of the DDP price will be paid by the tenderer. The purchaser will make the payment of DDP price after receipt of goods at consignee’s site in good condition as per payment terms

in the contract.

5.The prices quoted in foreign currency in column (e) shall be converted in Rupees at the selling rate of exchange applicable on the date of tender opening .The customs duty amount so worked out as

percentage of net CIP value in rupees will be taken for evaluation and comparison of tenders

Name_____________________________

Business address____________________

Place:________________________________________                                                                                                            Signature of Tenderer___________________

Date:_____________________________________                                                                                         Seal of Tenderer________________________

SECTION – XI PRICE SCHEDULE

PRICE SCHEDULE FOR COMPREHENSIVE MAINTENANCE CONTRACT AFTER WARRANTY PERIOD TO BE QUOTED FOR 3 YEARS

& ALSO FOR  5 YEARS CMC AND AMC TO BE QUOTED IN THE SAME FORMAT SUBSEQUENT TO 3 YEARS CMC ( THE PRICE

SUBSEQUENTLY QUOTED FOR AMC AND CMC WILL NOT BE CONSIDERED FOR EVALUATION PURPOSE)

| 1 | 2 | 3 | 4 |  |  |  |  | 5 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Item SI. No. | BRIEF DESCRIPTION OF GOODS | QUANTITY. (Nos.) | Comprehensive Maintenance Contract Cost for Each Unit year wise*. |  |  |  |  | Total Comprehensive Maintenance Contract Cost for 5 (or as specified) Years [3 x (4a+4b+4c+4d+4e)] |
|  |  |  | 1st | 2nd | 3rd | 4th | 5th |  |
|  |  |  | a | b | c | d | e |  |

  - * After completion of Warranty period

NOTE:-

- The cost of Comprehensive Maintenance Contract (CMC) which includes preventive maintenance including testing & calibration as per technical/ service /operational manual, labour and spares, after satisfactory completion of Warranty period may be quoted for next 5 years on yearly basis for complete equipment and Turnkey (if any).

- The cost of CMC may be quoted along with taxes applicable on the date of Tender Opening. The taxes to be paid extra, to be specifically stated. In the absence of any such stipulation the price will be taken inclusive of such taxes and no claim for the same will be entertained later.

- Cost of CMC will be added for Ranking/Evaluation purpose.

- The payment of CMC will be made as per clause GCC clause 21.1 (D).

- All software updates should be provided free of cost during CMC period.

- The supplier shall keep sufficient stock of spares required during Comprehensive Maintenance Contract period. In case the spares are required to be imported, it would be the responsibility of the supplier to import and get them custom cleared and pay all necessary duties.

Name________________________

Business Address________________________

Place: ___________________________ 	Signature of Tenderer________________________

Date: _________________________ 	Seal of the Tenderer________________________

###### SECTION – XI PRICE SCHEDULE

D)	PRICE SCHEDULE FOR TURNKEY

| Schedule No. | BRIEF TURNKEY DESCRIPTION OF GOODS | CONSIGNEE CODE | Turnkey price |
| --- | --- | --- | --- |

Note: -

- The cost of Turnkey as per Technical Specification (Section VII) may be quoted on lump sum along with taxes applicable on the date of Tender Opening. The taxes to be paid extra, to be specifically stated. In the absence of any such stipulation the price will be taken inclusive of such taxes and no claim for the same will be entertained later.

- Cost of Turnkey will be added for Ranking/Evaluation purpose.

- The payment of Turnkey will be made as per clause GCC clause 21.1 (c).

- The stipulations in Technical Specification will supersede above provisions

Name________________________

Business Address________________________

Place: ___________________________ 	Signature of Tenderer________________________

Date: _________________________ 	Seal of the Tenderer________________________

SECTION – XII

QUESTIONNAIRE

Fill up the Section XX – Check List for Tenderers and enclose with the Tender

- The tenderer should furnish specific answers to all the questions/issues mentioned in the Checklist. In case a question/issue does not apply to a tenderer, the same should be answered with the remark “not applicable”

- Wherever necessary and applicable, the tenderer shall enclose certified copy as documentary proof/ evidence to substantiate the corresponding statement.

- In case a tenderer furnishes a wrong or evasive answer against any of the question/issues mentioned in the Checklist, its tender will be liable to be ignored.

SECTION – XIII

BANK GUARANTEE FORM FOR EMD

Whereas ______________________________ (hereinafter called the “Tenderer”) has submitted its quotation dated ___________________ for the supply of _____________________________ (hereinafter called the “tender”) against the purchaser’s tender enquiry No. __________________________________ Know all persons by these presents that we ______________________________ of _______________________________ (Hereinafter called the “Bank”) having our registered office at ________________________________ are bound unto ____________________________ (hereinafter called the “Purchaser) in the sum of _________________________ for which payment will and truly to be made to the said Purchaser, the Bank binds itself, its successors and assigns by these presents. Sealed with the Common Seal of the said Bank this _____________________ day of _______ 20____. The conditions of this obligation are:

- (1) If the Tenderer withdraws or amends, impairs or derogates from the tender in any respect within the period of validity of this tender.

- (2) If the Tenderer having been notified of the acceptance of his tender by the Purchaser during the period of its validity:-

- a) fails or refuses to furnish the performance security for the due performance of the contract.

- or

- b) fails or refuses to accept/execute the contract.

- or

- c) if it comes to notice that the information/documents furnished in its tender is incorrect, false, misleading or forged

We undertake to pay the Purchaser up to the above amount upon receipt of its first written demand, without the Purchaser having to substantiate its demand, provided that in its demand the Purchaser will note that the amount claimed by it is due to it owing to the occurrence of one or both the two conditions, specifying the occurred condition(s).

This guarantee will remain in force for a period of forty-five days after the period of tender validity and any demand in respect thereof should reach the Bank not later than the above date.

(Signature of the authorised officer of the Bank)

Name and designation of the officer

Seal, name & address of the Bank and address of the Branch

SECTION – XIV

MANUFACTURER’S AUTHORISATION FORM

To

DVP(Tech.), HLL Lifecare Ltd, (A Government of India Enterprise), Infrastructure Development Division, “Adarsh” TC 6/1781, Vettamukku, Thirumala P.O, Thiruvananthapuram-695006.

Dear Sirs,

Ref. Your TE document No ____________, dated _____________

We, ___________________________________ who are proven and reputable manufacturers of___________________________(name and description of the goods offered in the tender) having factories at_______________________________, hereby authorise Messrs______________________________(name and address of the agent) to submit a tender, process the same further and enter into a contract with you against your requirement as contained in the above referred TE documents for the above goods manufactured by us.

We further confirm that no supplier or firm or individual other than Messrs. __________________________ (name and address of the above agent) is authorised to submit a tender, process the same further and enter into a contract with you against your requirement as contained in the above referred TE documents for the above goods manufactured by us. Agency agreement with them giving details of agency commission shall be provided.

We also hereby extend our full warranty, CMC as applicable as per clause 15 of the General Conditions of Contract, read with modification, if any, in the Special Conditions of Contract for the goods and services offered for supply by the above firm against this TE document.

We also hereby confirm that we would be responsible for the satisfactory execution of contract placed on the authorised agent.

We also confirm that the price quoted by our agent shall not exceed than that which we would have quoted directly.

Yours faithfully,

[Signature with date, name and designation]

for and on behalf of Messrs___________________________

[Name & address of the manufacturers]

Note : 1.	This letter of authorisation should be on the letter head of the manufacturing firm and should be signed by a person competent and having the power of attorney to legally bind the manufacturer.

2. Original letter may be sent.

SECTION – XV

BANK GUARANTEE FORM FOR PERFORMANCE SECURITY/ CMC SECURITY

To

Head of Hospital/Institute/Medical College

WHEREAS _____________________________ (Name and address of the supplier) (Hereinafter called “the supplier”) has undertaken, in pursuance of contract no________________________ dated _____________ to supply (description of goods and services) (herein after called “the contract”).

AND WHEREAS it has been stipulated by you in the said contract that the supplier shall furnish you with a bank guarantee by a scheduled commercial bank recognised by you for the sum specified therein as security for compliance with its obligations in accordance with the contract;

AND WHEREAS we have agreed to give the supplier such a bank guarantee;

NOW THEREFORE we hereby affirm that we are guarantors and responsible to you, on behalf of the supplier, up to a total of. ________________________ (Amount of the guarantee in words and figures), and we undertake to pay you, upon your first written demand declaring the supplier to be in default under the contract and without cavil or argument, any sum or sums within the limits of (amount of guarantee) as aforesaid, without your needing to prove or to show grounds or reasons for your demand or the sum specified therein.

We hereby waive the necessity of your demanding the said debt from the supplier before presenting us with the demand.

We further agree that no change or addition to or other modification of the terms of the contract to be performed there under or of any of the contract documents which may be made between you and the supplier shall in any way release us from any liability under this guarantee and we hereby waive notice of any such change, addition or modification.

This guarantee shall be valid up to 30 (thirty) months from the date of Notification of Award i.e  up to ----------- (indicate date)

…………………………….

(Signature with date of the authorised officer of the Bank)

………………………………………………………….

Name and designation of the officer

………………………………………………………….

………………………………………………………….

Seal, name & address of the Bank and address of the Branch

SECTION – XVI

CONTRACT FORM - A

CONTRACT FORM FOR SUPPLY, INSTALLATION, COMMISSIONING, HANDING OVER, TRIAL RUN, TRAINING OF OPERATORS & WARRANTY OF GOODS

(Address of the Purchaser’s/Consignee’s

office issuing the contract)

Contract No___________ dated______________

This is in continuation to this office’s Notification of Award No_______ dated ______

1. Name & address of the Supplier: ______________________________

2. Purchaser’s TE document No________ dated____________ and subsequent Amendment No____________, dated_________ (if any), issued by the purchaser

3. Supplier’s Tender No_________ dated__________ and subsequent communication(s) No____________ dated _________ (if any), exchanged between the supplier and the purchaser in connection with this tender.

4. In addition to this Contract Form, the following documents etc, which are included in the documents mentioned under paragraphs 2 and 3 above, shall also be deemed to form and be read and construed as integral part of this contract:

  - (i) General Conditions of Contract;

  - (ii) Special Conditions of Contract;

- (iii) List of Requirements;

- (iv) Technical Specifications;

- (v) Quality Control Requirements;

- (vi) Tender Form furnished by the supplier;

- (vii) Price Schedule(s) furnished by the supplier in its tender;

- (viii) Manufacturers’ Authorisation Form (if applicable for this tender);

- (ix) Purchaser’s Notification of Award

Note : The words and expressions used in this contract shall have the same meanings as are respectively assigned to them in the conditions of contract referred to above. Further, the definitions and abbreviations incorporated under clause 1 of Section II – ‘General Instructions to Tenderers’ of the Purchaser’s TE document shall also apply to this contract.

5. Some terms, conditions, stipulations etc. out of the above-referred documents are reproduced below for ready reference:

(i) Brief particulars of the goods and services which shall be supplied/ provided by the supplier are as under:

| Schedule No. | Brief description of goods/services | Accounting unit | Quantity to be supplied | Unit Price | Total price | Terms of delivery |
| --- | --- | --- | --- | --- | --- | --- |

Any other additional services (if applicable) and cost thereof: _____________________

Total value (in figure) ____________ (In words) ___________________________

- Delivery schedule

(iii) Details of Performance Security

(iv) Quality Control

(a) Mode(s), stage(s) and place(s) of conducting inspections and tests.

(b) Designation and address of purchaser’s inspecting officer

(v) Destination and despatch instructions

(vi) Consignee, including port consignee, if any

- Warranty clause

- Payment terms

- Paying authority

____________________________

(Signature, name and address

of the Purchaser’s/Consignee’s authorised official)

For and on behalf of__________________

Received and accepted this contract

(Signature, name and address of the supplier’s executive

duly authorised to sign on behalf of the supplier)

For and on behalf of _________________________

(Name and address of the supplier)

(Seal of the supplier)

Date: _________________________

Place: _________________________

SECTION – XVI

CONTRACT FORM – B

CONTRACT FORM FOR ANNUAL COMPREHENSIVE MAINTENANCE CONTRACT

Annual CM Contract No._______________________	dated_________________

Between

(Address of Head of Hospital/Institute/Medical College)

And

(Name & Address of the Supplier)

Ref:	Contract No___________ dated______________ (Contract No. & date of Contract for supply, installation, commissioning, handing over, Trial run, Training of operators & warranty of goods)

In continuation to the above referred contract

- The Contract of Annual Comprehensive Maintenance is hereby concluded as under: -

| 1 | 2 | 3 | 4 |  |  |  |  | 5 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Schedule No. | BRIEF DESCRIPTION OF GOODS | QUANTITY. (Nos.) | Annual Comprehensive Maintenance Contract Cost for Each Unit year wise*. |  |  |  |  | Total Annual Comprehensive Maintenance Contract Cost for 5 Years [3 x (4a+4b+4c+4d+4e)] |
|  |  |  | 1st | 2nd | 3rd | 4th | 5th |  |
|  |  |  | a | b | c | d | e |  |

Total value (in figure) ____________ (In words) ___________________________

b)	The CMC commence from the date of expiry of all obligations under Warranty i.e. from______________ (date of expiry of Warranty) and will expire on ______________ (date of expiry of CMC)

c)	The cost of Annual Comprehensive Maintenance Contract (CMC) which includes preventive maintenance, labour and spares, after satisfactory completion of Warranty period may be quoted for next 5 years as contained in the above referred contract on yearly basis for complete equipment (including X ray tubes, Helium for MRI, Batteries for UPS, other vacummatic parts, _____ & _____) and Turnkey (if any).

d)	There will be 98% uptime warranty during CMC period on 24 (hrs) X 7 (days) X 365 (days) basis, with penalty, to extend CMC period by double the downtime period.

e)	During CMC period, the supplier shall visit at each consignee’s site for preventive maintenance including testing and calibration as per the manufacturer’s service/ technical/ operational manual. The supplier shall visit each consignee site as recommended in the manufacturer’s manual, but at least once in 6 months commencing from the date of the successful completion of warranty period for preventive maintenance of the goods.

f)	All software updates should be provided free of cost during CMC.

g)	The bank guarantee valid till ______________ [(fill the date) 2 months after expiry of entire CMC period] for an amount of Rs. _______________ [(fill amount) equivalent to 2.5 % of the cost of the equipment as per contract] shall be furnished in the prescribed format given in Section XV of the TE document, along with the signed copy of Annual CMC within a period of 21 (twenty one) days of issue of Annual CMC failing which the proceeds of Performance Security shall be payable to the Purchaser/Consignee.

h)	If there is any lapse in the performance of the CMC as per contract, the proceeds Annual CMC bank guarantee for an amount of Rs. __________ (equivalent to 2.5 % of the cost of the equipment as per contract) shall be payable to the Consignee.

i)	Payment terms:	The payment of Annual CMC will be made against the bills raised to the consignee by the supplier on six monthly basis after satisfactory completion of said period, duly certified by the Consignee. The payment will be made in Indian Rupees.

j)	Paying authority:	 ______________________ (name of the consignee i.e. authorised official)

____________________________

(Signature, name and address

of Institute official)

For and on behalf of__________________

Received and accepted this contract

(Signature, name and address of the supplier’s executive

duly authorised to sign on behalf of the supplier)

For and on behalf of _________________________

(Name and address of the supplier)

(Seal of the supplier)

Date: _________________________

Place: _________________________

SECTION – XVII

CONSIGNEE RECEIPT CERTIFICATE

(To be given by consignee’s authorized representative)

The following store (s) has/have been received in good condition:

- Contract No. & date	:______________________________

LC No: & date ( for LC shipments)     :______________________________

- Supplier’s Name	:______________________________

- Consignee’s Name & Address with telephone No. & Fax No. 	: ______________________________

- Name of the item supplied	:______________________________

- Quantity Supplied	:______________________________

- Date of Receipt by the Consignee	:______________________________

- Name and designation of Authorized Representative of Consignee	:______________________________

- Signature of Authorized Representative of Consignee with date	:______________________________

- Seal of the Consignee	:______________________________

SECTION – XVIII

Proforma of Final Acceptance Certificate by the Consignee

## No_______________

## Date_______________

#### To

M/s	_______________________

_______________________

_______________________

Subject:	Certificate of commissioning of equipment/plant.

This is to certify that the equipment(s)/plant(s) as detailed below has/have been received in good conditions along with all the standard and special accessories and a set of spares (subject to remarks in Para no.02) in accordance with the contract/technical specifications.  The same has been installed and commissioned.

(a)	Contract No______________________________________ dated__________________

(b)	Description of the equipment(s)/plants:_____________________________________

(c)	Equipment(s)/ plant(s) nos.:_______________________________________________

(d)	Quantity:________________________________________________________________

(e)	Bill of Loading/Air Way Bill/Railway

Receipt/ Goods Consignment Note no_______________ dated _________________

(f)	Name of the vessel/Transporter:___________________________________________

(g)	Name of the Consignee:___________________________________________________

(h)	Date of commissioning and proving test:____________________________________

Details of accessories/spares not yet supplied and recoveries to be made on that account.

Sl.	Description of Item	Quantity	Amount to be recovered No.

No.

The proving test has been done to our entire satisfaction and operators have been trained to operate the equipment(s)/plant(s).

The supplier has fulfilled its contractual obligations satisfactorily   ## or

The supplier has failed to fulfil its contractual obligations with regard to the following:

He has not adhered to the time schedule specified in the contract in dispatching the documents/drawings pursuant to ‘Technical Specifications’.

He has not supervised the commissioning of the equipment(s)/plant(s)in time, i.e. within the period specified in the contract from date of intimation by the Purchaser/Consignee in respect of the installation of the equipment(s)/plant(s).

The supplier as specified in the contract has not done training of personnel.

The extent of delay for each of the activities to be performed by the supplier in terms of the contract is

The amount of recovery on account of non-supply of accessories and spares is given under Para no.02.

The amount of recovery on account of failure of the supplier to meet his contractual obligations is__________________ (here indicate the amount).

Signature

Name

Designation with stamp

## Explanatory notes for filling up the certificate:

He has adhered to the time schedule specified in the contract in dispatching the documents/drawings pursuant to ‘Technical Specification’.

He has supervised the commissioning of the equipment(s)/plant(s) in time, i.e. within the time specified in the contract from date of intimation by the Purchaser/Consignee in respect of the installation of the equipment(s)/plant(s).

#### Training of personnel has been done by the supplier as specified in the contract

In the event of documents/drawings having not been supplied or installation and commissioning of the equipment(s)/plant(s) having been delayed on account of the supplier, the extent of delay should always be mentioned in clear terms.

SECTION – XIX

ANNEXURES

Annexure 1

DETAILS OF SHIPPING ARRANGEMENT FOR LINER CARGOES IN RESPECT OF

C & F/CIF/TURNKEY/F.O.R CONTRACTS FOR IMPORTS

- (a) SHIPMENT FROM PORTS OF U.K INCLUDING NORTHERN IRELAND (ALSO EIRE), FROM THE NORTH CONTINENT OF EUROPE (GERMANY, HOLLAND, BELGIUM, FRANCE, NORWAY, SWEDEN, DENMARK, FINLAND AND PORTS ON THE CONTINENTAL SEABOARD OF MEDITERRANIAN (I.E. FRENCH WESTERN ITALIAN PORTS), TO PORTS IN INDIA.

The Seller should arrange shipment of the goods by vessels belonging to the member lines of the India-Pakistan-Bangladesh Conference. If the Seller finds that the space on the ‘Conference Lines’ vessels is not available for any specific shipment, he should take up with India-Pakistan-Bangladesh Conference. Conferity House, East Grinstead, Sussex (UK), for providing shipping space and also inform the Shipping Co-ordination Officer, Ministry of Surface Transport, New Delhi, (Cable: TRANSCHART, NEW DELHI, Telex: VAHAN IN – 031 – 61157, 61158, 61159)

The Seller should arrange shipment through the Government of India’s Forwarding Agents, M/s Schenker & Co., 2000-Hamburg (Cable: SCHENKER CO., HAMBURG) OR obtain a certificate from them to the effect that shipment has been arranged in accordance with instructions of the Ministry of Surface Transport, (TRANSCHART), New Delhi.

(b)	SHIPMENT FORM PORTS OF U.K. INCLUDING NORTHERN

Goods under this contract would be shipped by the national shipping companies of the Contracting Parties operating bilateral shipping service and vessels under the flag of third countries in accordance with the Agreement between the Government of German Democratic Republic and the Government of the Republic of India in the Field of Merchant Shipping signed on 9.1.1979, as amended up-to-date.

- ISHIPMENT FROM ADRIATIC PORTS OF EASTERN ITALY AND YUGOSLAVIA

The seller should arrange shipment of the goods by vessels belonging to the following Indian member lines;

- The Shipping Purchaser of India Ltd.

- The Scindia Steam Navigation Co., Ltd

- India Steamship Co., Ltd

For the purpose of ascertaining the availability of suitable Indian vessels and granting dispensation in the event of their non-availability, the Seller should give adequate notice about the readiness of each consignment from time to time at least six weeks in advance of the required position to M/s Schenker & Co. 2000 HAMBURG (Cable: SCHENKER CO., HAMBURG) and also endorse a copy thereof to the Shipping Co-ordination Officer, Ministry of Surface Transport, New Delhi, (Cable: TRANSCHART, NEW DELHI, Telex: VAHAN IN – 031 – 61157, 61158, 61159)

The seller should arrange shipment through the Government of India’s Forwarding Agents M/s Schenker & Co. 2000 HAMBURG (Cable: SCHENKER CO., HAMBURG) or obtain certificate from them to the effect that shipment has been arranged in accordance with the instructions of the Ministry of Surface Transport, (TRANSCHART), New Delhi.

- SHIPMENT FROM POLAND & CZECHOSLOVAKIA

- IMPORTS FROM POLAND

Shipment under this contract would be made by the National flag lines of the two parties and vessels of the third flag conference lines, in accordance with the agreement between the Govt. of the Republic of India and the Govt. of the Polish People’s Republic regarding Shipping Co-operation dated 27.6.1960 as amended up-to-date.

- IMPORTS FROM CZECHOSLOVAKIA

Goods under this contract would be signed by the National flag lines of the two parties and vessels of the third flag conference lines, in accordance with the Agreement Co-operation in shipping between India and Czechoslovakia signed on 3.11.1978 and ratified on 19.12.1979, as amended up-to-date.

Shipping arrangement should be made by the Sellers in consultation with Resident Representative of the Indian Shipping Lines in Gdynia, Co., Morska Agencja W. Gdyniul, Pulaskiego 8, P.O. Box 246, Gdynia (Poland) – Telex : MG PL. 054301, Tel.: 207621, to whom details regarding contract number, nature of  cargo , quantity, port of lading, discharging, name of Government consignee, expected date of readiness of each consignment etc. should be furnish at least six weeks in advance of the required position, with a copy thereof endorsed to the Shipping Co-ordination Officer, Ministry of Surface Transport, (Chartering Wing), New Delhi, (Cable: TRANSCHART, NEW DELHI, Telex: VAHAN IN – 031 – 61157, 61158, 61159)

(e)	SHIPMENT FROM U.S.S.R

Shipment under this contract should be made in accordance with the agreement between the Government of the Republic of India and the Government of U.S.S.R on  Merchant Shipping 1976, as amended up-to-date, by vessels of Indo-Soviet shipping Service.

- SHIPMENT FROM JAPAN

The shipment of goods should be made of India vessels to the maximum extent possible subject to the minimum of 50%.

The Seller should arrange shipment of the goods in consultation with the Embassy of India in Japan, Tokyo to whom details regarding contract number, nature of cargo, quantity, port of loading/discharge, name of Govt. consignee, expected date of readiness of each consignment etc. should be furnished at least six weeks in advance of the required position.

Note:  The copies of such contracts are to be endorsed both to the Attached (commercial) embassy of India in Japan, Tokyo, and the shipping Co-ordination Officer, Ministry of Surface Transport, New Delhi.

(g)	SHIPMENT FROM AUSTRALIA, ALGERIA, BULGARIA, ROMANIA, EGYPY

The Seller shall arrange shipment of the goods by Indian flag vessels to the maximum extent possible subject to a minimum of 50 %. For the purpose of ascertaining the availability of suitable Indian vessels, the seller shall give adequate notice of not less than six weeks about the readiness of each consignment to the Shipping Purchaser of India Ltd., SHIPPING HOUSE, 245, Madame Cama Road, Bombay – 400 021 (CABLE: SHIPINDIA BOMBAY) and also endorse a copy thereof to the Shipping Co-ordination Officer, Ministry of Surface Transport, New Delhi, (Cable: TRANSCHART, NEW DELHI, Telex: VAHAN IN – 031 – 61157, 61158, 61159)

- SHIPMENT FROM PAKISTAN

The shipment of cargoes should be made by Indian vessels to the maximum extent possible subject to a minimum of 50 %.

Shipment arrangement should be made by the sellers in consultation with M/s Mogul Line Ltd., 16-Bank Street, Fort, Bombay – 400023 (Cable: MOGUL BOMBAY: Telex: 011 – 4049 MOGUL), to whom, details regarding contract number, nature of cargo, quantity, port of lading discharging, name of government consignee, expected date of readiness of each consignment etc. should be furnish at least six weeks in advance of the required position, with a copy thereof endorsed to the Shipping Co-ordination Officer, Ministry of Surface Transport, New Delhi, (Cable: TRANSCHART, NEW DELHI, Telex: VAHAN IN – 031 – 61157, 61158, 61159)

(i)	SHIPMENT FROM U.S ATLANTIC & GULF PORTS

The Seller should arrange shipment of the goods by vessels belonging to the member lines of the India – Pakistan – Bangladesh – Ceylon and Burma Outward Freight Conference. If the Seller finds that the space of the ‘Conference Lines’ vessels is not available for any specific shipment he should take up with India – Pakistan- Bangladesh – Ceylon and Burma Outward Freight Conference, 19, Rector Street, New York, N.Y. 10006 USA, for providing shipping space and also inform the Shipping Co-ordination Officer, Ministry of Surface Transport, New Delhi, (Cable: TRANSCHART, NEW DELHI, Telex: VAHAN IN – 031 – 61157, 61158, 61159)

- SHIPMENT FROM ST. LAWRENCE AN EASTERN CANADIAN PORTS

The Seller should arrange shipment of the goods by vessels belonging to the following shipping lines;

- The shipping Purchaser of India Ltd.

- The Scindia Steam Navigation Co., Ltd

If the Seller finds that the space in the vessels of these Lines is not available for any particular consignments, he should inform the Shipping Co-ordination Officer, Ministry of Surface Transport, New Delhi, (Cable: TRANSCHART, NEW DELHI, Telex: VAHAN IN – 031 – 61157, 61158, 61159) immediately so that dispensation from the shipping lines concerned to use alternative lifting may be sought.

(k)	SHIPMENT FROM WEST COAST PORTS OF U.S.S CANADA AND OTHER AREAS NOT SPECIFICALLY MENTIONED ABOVE

The Seller should arrange shipment of the goods by Indian vessels to the maximum extent possible subject to a minimum of 50 %. For the purpose of ascertaining the availability of suitable Indian vessels and granting dispensation in the event of their non-availability, the Seller should furnish the details regarding contract number, nature of cargo, quantity, port of lading, discharging, name of government consignee, expected date of readiness of each consignment etc. to the Shipping Co-ordination Officer, Ministry of Surface Transport, New Delhi, (Cable: TRANSCHART, NEW DELHI, Telex: VAHAN IN – 031 – 61157, 61158, 61159) at least six weeks in advance of the required position.

- BILLS OF LADING

  - C.I.F./C&F/TURNKEY SHIPMENTS

The Bills of lading should be drawn to indicate Shipper and ‘Consignee’ as under:

SHIPPER: 	The C.I.F (C&F)/TURNKEY SUPPLIERS concerned.

CONSIGNEE:	As per consignee’s particulars in the contract (The name an address of the ‘Port Consignee’ and ‘Ultimate’ both should be indicated).

  - F.O.R SHIPMENTS

The Bills of lading should be drawn to indicate shipper Consignee as under:

SHIPPER: 	The F.O.R suppliers Concerned

CONSIGNEE:	Supplier’s Indian Agent on order

#### Note:

- Moreover the name of the ‘Purchaser’ and ‘Ultimate’ Consignee should appear in the body of the Bills of Lading as the ‘Notify’ or as a remark.

- Two non-negotiable copies of the Bills of Lading indicating the freight amount and discount, if any allowed, should be forwarded to The Shipping Co-ordination Officer, Ministry of surface Transport (Chartering Wing), New Delhi after the shipment of each consignment is effected.

- The seller should avoid the use of over-aged vessels for the shipment of the goods under the contract and if so used the cost of additional. Insurance, if any, shall be borne by the seller.

SECTION – XX

CHECKLIST

### Name of Tenderer:

### Name of Manufacturer:

| Sl No. | Activity | Yes/ No/ NA | Page No. in the TE document | Remarks |
| --- | --- | --- | --- | --- |
| 1. a. | Have you enclosed EMD of required amount for the quoted schedules? |  |  |  |
| b. | In case EMD is furnished in the form of Bank Guarantee, has it been furnished as per Section XIII? |  |  |  |
| c. | In case Bank Guarantee is furnished, have you kept its validity of 165 days from Techno Commercial Tender Opening date as per clause 19 of GIT? |  |  |  |
| 2. a. | Have you enclosed duly filled Tender Form as per format in Section X? |  |  |  |
| b. | Have you enclosed Power of Attorney/Partnership Agreement in favour of the signatory attested by a Notary Public. |  |  |  |
| 3. | (a) Are you a SSI unit, registered with NSIC under Single point registration Scheme or registered with DGS&D for the quoted items ? If so, have you enclosed a copy of the registration certificate? (b) Are you enlisted with DGS&D as Indian Agent under the compulsory Enlistment Scheme of Ministry of Finance, Govt. of India? If so have you enclosed a copy of the enlistment certificate? |  |  |  |
| 4. a. | Have you enclosed clause-by-clause technical compliance statement for the quoted goods vis-à-vis the Technical specifications? |  |  |  |
| b. | In case of Technical deviations in the compliance statement, have you identified and marked the deviations? |  |  |  |
| 5. a. | Have you submitted satisfactory performance certificate from the end users in respect of all orders mentioned in the Proforma for performance statement in Sec. IX of TE document. |  |  |  |
| b. | Have you submitted copy of the supply order(s) and installation report? |  |  |  |
| 6. | (a) Have you submitted manufacturer’s authorization as per Section XIV? (b) Have you submitted a copy of the agreement between you and your Principal as per clause 14 of GIT? |  |  |  |
| 7. | (a) Have you submitted prices of goods, turnkey (if any), CMC etc. in the Price Bid as per Section XI? (b) Have you submitted with your Price Bid your Principal’s /Manufacturer’s Original proforma invoice indicating FOB value and Indian Agent Commission . |  |  |  |
| 8. | Have you kept validity of 120 days from the Techno Commercial Tender Opening date as per the TE document? |  |  |  |
| 9. a. | In case of Indian Tenderer, have you furnished Income Tax Account No. as allotted by the Income Tax Department of Government of India? |  |  |  |
| b. | In case of Foreign Tenderer, have you furnished Income Tax Account No. of your Indian Agent as allotted by the Income Tax Department of Government of India? |  |  |  |
| 10. | Have you intimated the name an full address of your Banker (s) along with your Account Number |  |  |  |
| 11. | (a) Have you fully accepted payment terms as per TE document? (b) Have you accepted “terms of delivery” as per TE document? |  |  |  |
| 12. | Have you fully accepted delivery period as per TE document? |  |  |  |
| 13. | Have you confirmed that the terms of delivery shall be “Delivery at Consignee Site” ? |  |  |  |
| 14. | Have you accepted the warranty as per TE document ? |  |  |  |
| 15. | Have you accepted all other terms and conditions of TE document? |  |  |  |
| 16. | (a) Have you furnished documents establishing your eligibility & qualification criteria as per TE documents? (b) Have you given “write up” as asked for in Qualification Criteria (Section IX) under Note 2 ? |  |  |  |
| 17 | Have you furnished Annual Report (Balance Sheet and Profit & Loss Account) for last three years prior to the date of Tender opening? |  |  |  |
| 18 | Have you submitted the certificate of incorporation? |  |  |  |

N.B.

1.  All pages of the Tender should be page numbered and indexed.

2.  The Tenderer may go through the checklist and ensure that all the documents/confirmations listed above are enclosed in the tender and no column is left blank. If any column is not applicable, it may be filled up as NA.

3 . It is the responsibility of tenderer  to go through the TE document to ensure furnishing all required

documents in addition to above, if any. Non submission of the above information/documents

makes the bid liable to be ignored without any further reference to the Bidder.

(Signature with date)

(Full name, designation & address of the person duly authorised sign on behalf of the Tenderer)

For and on behalf of

(Name, address and stamp of the tendering firm)

- Section – XXI

- Consignee Detail

Tertiary Cancer Care Centre at Government Medical College, Kozhikode, Ministry of Health & Family Welfare. Govt of Kerala.

NB:	The consignee will ensure timely issue of Road permit, NMIC, CDEC, Octroi

Exemption Certificates & Entry Tax Exemption Certificates, wherever applicable, to the suppliers.