एचएलएललाइफकेयरलिमिटेड/HLL LIFECARE LIMITED

भारतसरकारकाउद्यम/A GOVT.OF INDIA ENTERPRISE

आककुलमफ़ैक्टरि /Akkulam Factory, श्रीकारियम.पी.ओ। /Sreekariyam. P.O,

तिरुवनंतपुरम-६९५०१७ / THIRUVANANTHAPURAM-695 017

PHONE NO:0471 – 2442641, 2445930,2445935;

EMAIL: materialsaft@lifecarehll.com

### INVITATION FOR BID

(TECHNICAL BID)

FOR THE SETTING UP OF EXTRUSION FACILITY

AND SUPPLY OF MEDICAL GRADE PVC SHEET

FOR BLOOD BAG PRODUCTION

(TENDER No: HLL/AFT/MTLS/EXTRUSION FACILITY/2022-23 Dated 10-01-2023)

AT

AKKULAM FACTORY

THIRUVANANTHAPURAM

## INDEX

| Section | Topic | Page No. |
| --- | --- | --- |
| I | Notice Inviting Invitation For Bid (IFB) | 4-6 |
| II | Prequalification Criteria | 7-8 |
| II.A | Questionnaire for Prequalification Criteria | 9-10 |
| II.B | General Information of the Bidder | 11 |
| III | Technical Specification | 12-15 |
| IV | Terms & Conditions | 16-22 |
| V | Special Conditions to the Bidders | 23-29 |
| VI | Bid Security Declaration Form | 30 |
| VII | Indemnity Clause | 31 |
| VIII | Declaration | 32 |
| IX | Acceptance Letter | 33 |
| X | Pre-contract Integrity pact | 33-41 |
| XI | Price Bid Format | Annexure I |

एचएलएललाइफकेयरलिमिटेड/HLL LIFECARE LIMITED

भारतसरकारकाउद्यम/A GOVT.OF INDIA ENTERPRISE

आककुलमफ़ैक्टरि /Akkulam Factory, श्रीकारियम.पी.ओ। /Sreekariyam. P.O,

तिरुवनंतपुरम-६९५०१७ / THIRUVANANTHAPURAM-695 017

PHONE NO:0471 – 2442641, 2445930,2445935;

## EMAIL: materialsaft@lifecarehll.com

## CIN No U25193KL1966GOI002621

TENDER No: HLL/AFT/MTLS/EXTRUSION FACILITY/2022-23 			     10-01-2023

## SECTION I

## NOTICE INVITING INVITATION FOR BID (IFB)

HLL Lifecare Limited (HLL) is a Public Sector Enterprise under the Ministry of Health & Family Welfare, Government of India. Akkulam Factory, Thiruvananthapuram (AFT) is one of the manufacturing facilities of HLL Lifecare Limited and engaged in the production of medical devices and contraceptive products. Blood Bag is one of the products being manufactured at Akkulam Factory for the last more than 25 years. Medical Grade PVC Sheet is the main base material used for the production of Blood Bag. HLL-AFT is having one Sheet Extrusion Facility for producing PVC Sheet. Specially formulated Medical Grade PVC Pellets is the Raw material of PVC Sheet. In house Extrusion Facility is capable of catering only 60% of the requirement of the factory capacity.

Hence HLL, Akkulam Factory, Thiruvananthapuram invites Bids for Setting Up of Extrusion Facility for the Supply of Medical Grade PVC Sheet for the production of Blood Bag. Details of the bid is given below.

| Sl No | Particulars | Description |
| --- | --- | --- |
| 1 | Name of Item/Work | Invitation for Bid for Setting Up of Extrusion Facility and Supply of Medical Grade PVC Sheet for the production of Blood Bag |
| 2 | Period of Contract | A minimum period of 5 years from the date of commencement of Contract and further extendable on mutual consent basis. |
| 3 | Technical Requirement | Section VI |
| 4 | Bid Security/EMD | Rs 2,00,000/- |
| 5 | Eligibility criteria for Bidders | As per IFB, Prequalification Criteria |
| 6 | Date of Publishing Invitation for Bid (IFB) | 10-01-2023 |
| 7 | Date of Pre bid meeting (Online) | 18-01-2023 at 11:00 Hrs Venue: HLL Lifecare Limited, Akkulam Factory, Sreekariyam (PO), Thiruvananthapuram – 17 (VC also will be available) |
| 8 | Last Date for Online submission of Bids | 31-01-2023 |
| 9 | Date of opening of Bid | 31-01-2023 |
| 10 | HLL A/c Details for payment of EMD (Payment mode: NEFT/RTGS) | Name of Bank : State Bank of India A/c number : 30307559515 IFSC Code : SBIN0004350 SWIFT Code : SBININBB564 Branch name : Commercial Branch, Branch : Thycaud, Thiruvananthapuram. |

- This IFB is an open Tender and is being published in public domain, Government e-Market Place (GeM Portal). Amendment/Corrigendum/Addendum with respect to this tender will be published through GeM portal & in our website, www.lifecarehll.com. Bidder should take into account any Amendment/Corrigendum/Addendum published for this IFB before submitting their bids. Bidders have to submit all supporting documents to against the Prequalification Criteria for eligibility, as required in this IFB document along with the bid.

- EARNEST MONEY DEPOSIT (EMD)

The bidder shall furnish, as part of this bid, Earnest Money Deposit (EMD) for an amount of Rs.2,00,000/-. The Earnest Money Deposit (Bid Security) is required to protect the purchaser against the risk of bidder’s conduct, which would warrant the forfeiture of bid security. The bid security shall be paid through RTGS/NEFT transfer.

The bid security of the unsuccessful bidder will be discharged / returned as promptly as possible, but not later than 30 days after the expiry of the period of the bid validity.

- MICRO/SMALL/Start-Up units interested in availing exemption from payment of bid security should submit a valid copy of their registration certificate issued by the concerned DIC or NSIC/Udyog Aadhaar. If the bidder is a MSE, it shall declare in the bid document the Udyog Aadhaar Memorandum Number issued to it under the MSME Act, 2006. If a MSE bidder do not furnish the UAM Number along with bid documents, such MSE unit will not be eligible for the benefits available under Public Procurement Policy for MSEs Order 2012. But the Party has to provide Performance Security/Security Deposit if Tender is awarded to them.

- NSIC /SSI/MSE units have to submit a Bid Security Declaration Form to avail Bid Security exemption. The offer shall be firm in Indian Rupees. No Foreign exchange will be made available by the purchaser. The supplier shall quote as per price schedule for all the items given in schedule of requirement.

- Preference to Make In India (MII) Manufacturers: Preference shall be given to Class I & II local Contractor as defined in Public Procurement (Preference to Make in India), Order 2017 as amended from time to time and its subsequent Orders/Notifications issued by concerned Nodal Ministry for specific Goods/Products. The minimum local content to qualify as a Class 1 local Contractor will be as defined in Public Procurement (Preference to Make in India), Order 2017.

- Bids submitted by the bidders will be evaluated based on the eligibility criteria and responsiveness to the requirements of HLL. The Price Bids of bidders who fulfill the Prequalification Criteria (Technical Bid Requirements) only will be opened for further evaluation. HLL Lifecare Limited reserves the right to reject any or all of the bids without assigning any reason.

- In the event of the date specified for bid receipt and opening being declared as a closed holiday for HLL’s office, the due date for submission of bids and opening of bids will be the following working day at the appointed times.

- Bidders shall submit a written Power of Attorney authorizing the signatory/ Signatories of the tender to commit the tender.

- The complete set of documents as required as per the ATC should be submitted / uploaded in the GeM Portal along with Bid Security (Earnest Money Deposit- EMD). Any bid not accompanied by Bid Security will be rejected.

- Bids will be opened in the GeM Portal on the specified date and time, at the office of Joint General Manager, Materials Department, HLL Lifecare Limited, Akkulam Factory, Sreekaryam PO, Thriruvananthapuram, Kerala, India- 695017.

- Cancellation or creation of a document such as Power of Attorney, Partnership deed, Constitution of firm etc., which may have bearing on the Tender/Contract shall be communicated forthwith in writing / through GeM portal by the Tenderer to HLL.

- HLL may, at its discretion, extend this deadline for submission of bids by amending the Bid Documents for any other reasons, in which case all rights and obligations of the HLL and Bidders previously subject to the deadline will thereafter be subject to the deadline as extended. Bids sent by Postal/Courier/Email/Fax/Telegraph will not be accepted.

- The company reserves the right to change the qualifying criteria at their discretion and to cancel the tender or reject any of the bids without assigning any reason thereof.

- Any changes regarding extension of tender opening dates in GeM Portal shall be communicated through our website.

- HLL Lifecare Limited reserves the right to reject any or all of the offers without assigning any reason.

- Unit Chief, HLL Lifecare Limited, Akkulam Factory reserves right to cancel this tender at any point of time.

Joint General Manager (Materials)

एचएलएललाइफकेयरलिमिटेड/HLL LIFECARE LIMITED

भारतसरकारकाउद्यम/A GOVT.OF INDIA ENTERPRISE

आककुलमफ़ैक्टरि /Akkulam Factory, श्रीकारियम.पी.ओ। /Sreekariyam. P.O,

तिरुवनंतपुरम-६९५०१७ / THIRUVANANTHAPURAM-695 017

TENDER No: HLL/AFT/MTLS/EXTRUSION FACILITY/2022-23 			     10-01-2023

## SECTION II

- PRE-QUALIFICATION CRITERIA OF THE BIDDER

- The Tenderer should either having Extrusion Facility / Calendaring Facility or be ready to set up an Extrusion Facility at his own expense at his own / rental premises for the production of Medical grade DEHP based PVC Sheet for Blood Bag Production and the Annual production capacity should be sufficient to meet the requirement of HLL.

- Vendor shall be ready to provide dedicated Sheet Extruder Machine / Calendaring Machine in Clean Room Facility for producing Medical Grade PVC Sheet. The Extrusion / Calendaring Machine should be capable of delivering the Medical Grade PVC Sheet as per the specification and standards mentioned in the IFB.

- The average annual turnover of the bidder for the last 3 years ending March 2022 should be Rs. 5 Crores. Vendors are required to submit documents relating to last three years in support of their claim. SSI/MSE/NSIC/Startup Units are eligible for exemption of Turnover, subject to the submission of copy of their valid MSME / Udyog Aadhaar registration certificate.

- The tenderer should have ISO 9001 certification and a copy of latest audit report of the certifying body should be attached along with the tender document OR should be ready to take the certification within 3 months from the date of award of contract/LOI/Signing Agreement and qualification of PVC Sheet.

- The tenderer should have ISO 13485 Certification for Medical Device Manufacturing and a copy of latest audit report of the certifying body should be attached along with the tender document OR should be ready to take the certification within 3 months from the date of award of contract/LOI/Signing Agreement and qualification of PVC Sheet.

- The tenderer should have ISO Class 8 Clean Room facility for the production or should be ready to set up Clean Room confirming ISO Class 8 compliance & certification within 3 months from the date of award of contract/LOI/Signing Agreement. The Clean Room must be tested and validated as per the regulatory guidelines, ISO 14644.

- The vendor should have In-house testing facility to check the product for HLL Parameters and should be able to issue test certificate for each batch/lot.

- EMD of Rs.2,00,000/- as per the tender conditions shall be paid separately, through RTGS/NEFT transfer. SSI/MSE/NSIC/Startup Units interested in availing exemption from payment of EMD should submit a valid copy of their MSME / Udyog Aadhaar registration certificate and also have to submit the Bid Security Declaration.

- The Successful tenderer should provide Security Deposit (SD) by DD / Bank Guarantee from Nationalized/Scheduled Bank before commencement of the work, if the work is awarded. Security Deposit will be 3 % of the order value to a maximum of Rs.2,00,000/-, which will recovered from the running bills of the supplies.

- The successful tenderer should be ready to execute an Agreement with HLL for a period of 5 Years for the supply of Medical grade DEHP based PVC Sheet in tune with HLL requirement at the finalized rate.

- A DECLARATION (as given in SCHEDULE-I) from the tenderer that the offer being submitted is as per the requirements given in this Bid is to be attached along with the Technical Bid.

Technical Evaluation team will be visiting the facility of the supplier to ensure and confirm the adherence of the Prequalification Criteria and to qualify the facility of the successful bidder before commencing the supplies.

Joint General Manager (Materials)

Questionnaire for Pre-Qualification (Minimum Eligibility Criteria)

SECTION II

| 1. | Do you have Extrusion / Calendaring Facility for the production of Medical grade DEHP based PVC Sheet | YES/NO |
| --- | --- | --- |
| 1.a | If Yes, what is the Annual Production Capacity & Details of the Machine (Make, Year of Manufacture, etc.) |  |
| 1.b | If No, Are you ready to set up Extrusion Facility at your own expense at your own / rental premises for the production of Medical grade DEHP based PVC Sheet? | YES/NO |
| 1.c | What is the Annual Production capacity planned? |  |
| 1.d | Are you planning to set up the facility in your Own Premises OR in Rental Premises for the production | Own Premises / Rental Premises |
| 2. | Do you have the facility to confirm the compliance of ISO Class 8 Clean Room & The Clean Room to be tested and validated as per the regulatory guidelines as per ISO 14644 ? | YES/NO |
| 2.a | If No, are you ready to set up the facility? | YES/NO |
| 3. | How much time is required to set up the facility and to start supplies as per HLL Pre-Qualification Criteria? |  |
| 4. | Have you attached Annual Turnover for the last three years? | YES/NO |
| a. | 2021 – 22 |  |
| b. | 2020 – 21 |  |
| c. | 2019 - 20 |  |
| d. | If No, Are you an MSE /Start UP Vendor | YES/NO |
| 5. | Do you have ISO 9001 certification? | YES/NO |
| 5.a | If Yes, Have you attached copies of latest audit report of the certifying body? | YES/NO |
| 5.b | If No, are you ready to take the certification? | YES/NO |
| 6. | Do you have ISO 13485 Certification for Medical Device Manufacturing? | YES/NO |
| 6.a | If Yes, Have you attached copies of latest audit report of the certifying body? | YES/NO |
| 6.b | If No, are you ready to take the certification? | YES/NO |
| 7. | Do you have testing facility to check HLL’s parameters? | YES/NO |
| 7.a | Do you have a Quality Control Department? If yes, Specify details | YES/NO |
| 7.b | Do you have inspection for Incoming Materials? If yes, Specify details | YES/NO |
| 7.c | Do you have In-Process Inspection Facilities? If yes, Specify details | YES/NO |
| 7.d | Do you have Final Inspection of the product? If yes, Specify details | YES/NO |
| 8. | Have you submitted Earnest Money Deposit ? | YES/NO |
| 8.a | Are you eligible for EMD Exemption | YES/NO |
| 8.b | Are you an MSE /Start UP Vendor ? | YES/NO |
| 8.c | If Yes, specify the social Status (General/Women/SC&ST/PH) |  |
| 8.d | If Yes, Have you submitted Bid Security Declaration Form? | YES/NO |
| 9. | Are you ready to submit the Security Deposit (SD) as prescribed in the Tender Document ? | YES/NO |
| 10. | Are you ready to execute an agreement with HLL for a period of 5 Years for the supply of Medical grade DEHP based PVC Sheet in tune with HLL requirement | YES/NO |
| 11. | Have you attached duly Signed & Sealed DECLARATION with your bid ? | YES/NO |
| 12. | Are you an existing vendor of HLL Lifecare Limited? If yes, Since when………………… | YES/NO |

All the information provided herein is True & Correct

PLACE:      						NAME & SIGNATURE OF THE APPLICANT

DATE:

(WITH OFFICE SEAL)

General information of the manufacturer / supplier

SECTION II

| Sl No | Description | Status |
| --- | --- | --- |
| 1 | Name & Address of the Bidder |  |
| a. | Telephone No. |  |
| b. | E-mail Address |  |
| c. | Name of contact person |  |
| d. | Whether proprietary/partnership/ Limited company. |  |
| e. | Specify whether SSI / MSE unit |  |
| f. | If Yes, please specify the social status of the entrepreneur (General/ SC&ST/ Women/ PH) |  |
| g. | Specify whether Startup unit |  |
| h. | GST NO |  |
| i. | PAN No |  |
| j. | HSN Code of the Item |  |
| 2. | Bank Details |  |
| a. | Name |  |
| b. | Branch & Address of your Banker |  |
| c. | Account No. |  |
| d. | Swift Code / IFSC Code |  |
| 3. | Details of your Production and Products |  |
| a. | How many years have you been in the business of the production? |  |
| 4. | What is your annual production capacity |  |
| 5. | Do you have ISO 9001 certification, If yes, please give the details. |  |
| 6. | Do you have an ISO 14001 certification, If yes, please give the details. |  |
| 7. | Do you have an ISO 45001 certification, If yes, please give the details. |  |
| 8. | Do you have any Subsidiary Units? If yes mention the details |  |
| 9. | Furnish the Name and address of the preferred Transporter through whom the material is transported to HLL AND rejected material is to be returned on freight to pay basis |  |

All the information provided herein is true & correct.

PLACE:                                                       NAME & SIGNATURE OF THE APPLICANT

DATE :

(WITH OFFICE SEAL)

NOTE: Forms must be completed in all respect and supported by relevant documents.

एचएलएललाइफकेयरलिमिटेड/HLL LIFECARE LIMITED

भारतसरकारकाउद्यम/A GOVT.OF INDIA ENTERPRISE

आककुलमफ़ैक्टरि /Akkulam Factory, श्रीकारियम.पी.ओ। /Sreekariyam. P.O,

तिरुवनंतपुरम-६९५०१७ / THIRUVANANTHAPURAM-695 017

TENDER No: HLL/AFT/MTLS/EXTRUSION FACILITY/2022-23 			     10-01-2023

## SECTION III

## TECHNICAL SPECIFICATION

- SCOPE OF INVITATION FOR BID (IFB)

HLL Lifecare Limited (HLL), a public sector undertaking under the Ministry of Health & Family Welfare, Government of India invites bids for Setting-Up of Extrusion/Calendaring Facility and Supply of Medical Grade PVC Sheet for the production of Blood Bag.

- REQUIREMENT OF CONTRACT

Successful Bidder shall set up a Sheet Extrusion / Calendaring Facility with for extruding Medical Grade PVC Sheet for the supply as per the specification mentioned in the IFB to HLL Life care Limited, Akkulam Factory, Thiruvananthapuram, Kerala for the production of Blood Bag. The Extrusion/Calendaring Facility shall be at his own expense and at his own / rental premises. The Annual production capacity of the Extrusion Facility to be set up should be sufficient to meet the requirement of HLL.

- SPECIFICATION OF THE MEDICAL GRADE PVC SHEET TO BE EXTRUDED

Specification of the Sheet to be extruded and supplied as detailed below.

- Item Name: Medical grade PVC sheet for Blood bag

- Raw Material of the Sheet: DEHP based Medical grade soft PVC granules (manufactured with PVC Resin having K-value 76 and other compounding ingredients as per HLL approved grades/sources and one side embossing. The raw materials should be procured only from HLL approved vendors, and the same is to be produced only as per HLL formulation. Re-processed material should not be used for the production of sheet, only virgin material is to be used for the production of PVC sheet for Blood bags. (very critical requirement).

- Width of Sheet (4 types): 200 mm, 220 mm, 240 mm, 260 mm (Tolerance +/- 1mm)

- Thickness of Sheet: 0.40mm +/- 0.02mm

- Embossing: One side embossed (Ra value:  12 +/- 3 micro meter and the other side shall be smooth (Ra value: 2.5 +/-1 micro meter)

- Sheet Shore-A Hardness: 75 to 80

- Sheet Tensile strength: Should be greater than or equal to 180 Kg/cm2

- Uniform thickness shall be maintained throughout the roll in the range of 0.40 mm +/- 0.02 mm. There shall not be any thickness variation in the transverse/machine direction throughout the roll.

- Recommended weight of one Roll sheet (Bobbin weight + Sheet weight): 20 Kg +/- 0.5 Kg and 200 to 225 Mtrs / Roll.

- Sheet Winding: Sheet shall be wounded over Bobbin with required tightness to suit the HLL HF Welding Machine.

- Straightness of sheet during unwinding: The sideways deflection (banana shape) shall not exceed 5 cm over an unwound length (distance) of 10 meters.

- ID of PVC Bobbin for sheet winding: 85 +/- 1mm, Length of bobbin shall be the same as sheet width.

- Shrinkage after steam sterilization (119Co for 30 minutes): Less than 3% - 5% of initial dimension on either direction (longitudinal and transverse)

- Visual parameters: The sheet must be free from all types of foreign particles, black spot, embedded dust, material degradation, air bubbles, colour change, pin holes, cracks, patches, flow marks/lines, irregular embossing, irregular side cutting / slitting, etc.

- Packaging & storage: Each roll of sheet must be packed in Polyethylene covers and the same will be packed in good quality Paper board cartons and then the cartons shall be stretch packed to prevent damage during transportation and external contamination.

- Identification shall be provided on each roll with Batch No according to HLL nomenclature.

- Test Certificates: Supplier should provide Inspection Report /Certificate of Analysis (COA) with details of raw material batch number and vendor details for each lot of sheet.

- Successful Bidder may source the materials from the qualified vendors of HLL for PVC Pellets to ensure the formulation of pellets and there by confirming the quality of sheet to be compatible with HLL High Frequency Welding Machine. HLL will share the details of the pellets supplier to the successful bidder. The cost of PVC Pellets shall be in the scope of successful bidder.

- PREFERED MANUFACTURING PROCESS OF SHEET

The preferred manufacturing process of the sheet shall be Extrusion / Calendaring Process. The Extrusion / Calendaring Machine should be capable of delivering the Medical Grade PVC Sheet as per the specification and standards mentioned in the IFB. The Extrusion facility should meet the quality requirements mentioned in the Prequalification Criteria. HLL Technical Evaluation team will be visiting the facility of the supplier to ensure and confirm the adherence of the Prequalification Criteria and to qualify the facility of the successful bidder before commencing the supplies.

- INPROCESS INSPECTION REQUIREMENTS

The PVC sheet should meet all the visual and dimensional requirements as per the specifications Frequency of inspection-

- Visual Inspection: continuous monitoring

- Dimensional Inspection: Production start-up and every 15 minutes

- Autoclaving / shrinkage test: Once in every shift

- QUALITY CONTROL SAMPLING ACCEPTANCE CRITERIA

The Quality Sampling shall be as per standard: ISO 2859-1

- For visual inspection:

Level II Double Normal

For major defects	: AQL: 0.40

For minor defects	: AQL: 4.00

b) For Dimensional Inspection

S-1, Single reduced.

For critical dimensions	: AQL: 0.40 Accept / Reject: 0/1

For non-critical dimensions	: AQL: 4.0

Major visual inspection parameters and its possible defects are given below.

| Item name | Critical Dimensions | Major Visual defects |
| --- | --- | --- |
| Medical grade PVC sheet for Blood bag | - Sheet Thickness: 0.40mm +/- 0.02mm - Shrinkage: 3% - 5% of initial dimension - Roughness (Embossed): 12 +/- 3 micro meter - Roughness (Embossed): Ra value: 2.5 +/-1 micro meter - Straightness of sheet during unwinding: The sideways deflection (banana shape) shall not exceed 5 cm over an unwound length (distance) of 10 meters. - Shrinkage after steam sterilization: Less than 3% - 5% of initial dimension on either direction | - Black spot, embedded dust, material degradation, air bubbles, pin holes, cracks, irregular embossing, irregular side cutting / slitting. - Loose / non-uniform winding of sheet on Bobbin - Deformation of sheet before and after autoclave / steam sterilization. |

- INPROCESS INSPECTION REQUIREMENTS

- Sheet QC Sampling and Acceptance criteria:

Pre shipment Inspection by the Supplier: Each batch of finished product must be inspected as per the acceptance criteria (As mentioned above Visual and dimensional inspection criteria) before dispatch of each consignment and should ensure that the lot fully complies with the quality requirements. Sample shall be taken on √n+1 basis. (n= no. of boxes / sheet rolls)

- Documentation requirements:

Sheet must be supplied along with the necessary production & QC documents contain the following details.

- Production document: Comprising date of production, Name of the operator, Product batch No, Total Qty produced, Accepted Qty, Rejected Qty. Type of rejection and Critical Process Parameters (like, Barrel, Head, Die, Take-up roll and Cooling roll Temperature, Screw speed and Take up roll speed).

- QC Document: Contains the details of in-process inspection & finished product pre-shipment inspection details(Batch No, Date, Time of sampling, Sample size, Product quality status, Instruments / Gauges used, Inspected by, Verified by, etc.) and Final Test Report.

- Packaging Details: Date, Batch No, Quantity, QC Passed sticker etc.

- Documentary proof for the usage of approved Raw material: Copy of raw material test certificate and copy of invoice from the raw material supplier shall be sent along with each consignment of sheet to ensure that, proper material is used for the production.

- The quantity mentioned below is the approximate requirement and HLL reserves the right to increase or decrease the quantity at any time during the contract period. The contractor is bound to execute all the supply orders placed during the contract period as per the requirement of HLL.

- INCOMING QA INSPECTION, TESTING & QUALIFICATION:

Qualification of the Medical Grade PVC Sheet extruded by the bidder shall undergo the Incoming QA Inspection and Testing as per the protocol of HLL-AFT. Upon qualification of the PVC Sheet by QA Department, same will be handed over to Production Department for further operational suitability.

- OPERATIONAL SUTABILITY:

Medical Grade PVC Sheet qualified by the QA Department will be taken for the trail run in the High Frequency Welding Machine for confirming the Performance & Operational Suitability in the plant. Performance & Operational qualification includes the rejection of bags due to the quality of sheet and machine acceptance of sheet.

- Technical Team of HLL and the external auditors shall conduct  periodic inspection / audit to the extrusion facility to ensure the quality and cGMP Standards.

- QUANTITY REQUIREMENT

| Sl.No. | Item Name | Approximate quantity requirement per Annum in “Tons (MT)” |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  | 1stYear | 2ndYear | 3rdYear | 4thYear | 5th Year |
| 1. | MEDICAL GRADE PVC SHEETFOR BLOOD BAG | 200 | 200 | 200 | 200 | 200 |

- The quantity mentioned above is the approximate requirement and HLL reserves the right to modify it at any time during the contract period. The contractor is bound to execute all the supply orders placed during the contract period.

- TIME LINE TO SET UP EXTRUSION FACILITY & START SUPPLY

The Time Line to set up the extrusion facility and start the trails, qualification and supply to HLL shall be a maximum of 6 Months from the date of Letter of Intent / Agreement. Successful bidder shall HLL submit a detailed schedule of activities, which confined to maximum period of 6 Months to start supply the PVC Sheet as per the specification of HLL. Representatives / the third party deployed by HLL will visit the facility to review the progress of setting up of the Extrusion Facility.

We hereby agree to supply the above materials as per the specification listed herein.

PLACE:                                   			NAME & SIGNATURE OF THE APPLICANT

DATE:

(WITH OFFICE SEAL)

एचएलएललाइफकेयरलिमिटेड/HLL LIFECARE LIMITED

भारतसरकारकाउद्यम/A GOVT.OF INDIA ENTERPRISE

आककुलमफ़ैक्टरि /Akkulam Factory, श्रीकारियम.पी.ओ। /Sreekariyam. P.O,

तिरुवनंतपुरम-६९५०१७ / THIRUVANANTHAPURAM-695 017

TENDER No: HLL/AFT/MTLS/EXTRUSION FACILITY/2022-23 			     10-01-2023

## SECTION IV

## TERMS & CONDITIONS

- The Tender should be completed in all respects.  Incomplete tenders are liable to be rejected. EMD shall be made as per Tender condition and the supporting documents shall be uploaded in the GeM Portal. Those parties who are exempted from EMD shall submit the supporting documents along with Bid Security Declaration in the Portal. SSI/MSE units interested in availing exemption from payment of bid security should submit a valid copy of their registration certificate issued by the concerned DIC or NSIC/Udyog Aadhaar.

- Bidder should submit all relevant documents to prove the Pre-Qualification Criteria in the GeM Portal. Bidder shall ensure that following documents shall be attached in GeM Portal along with the Bid without fail.

- Copies of the Quality Certification such as ISO 9001, ISO 13485, etc.

- Facility Certification such as  ISO 14644 for Class 8 Clean Room

- Duly signed & sealed Questionnaire for Pre-Qualification

- Duly signed & sealed General information of the manufacturer / supplier

- Earnest Money Deposit (EMD) Payment Details

- Copy of Udhyog Aadhaar in the case of MSE Bidder

- Duly signed & sealed Bid Security Form in the case of MSE Bidder

- Duly signed & sealed INDEMNITY, DECLARATION & ACCEPTANCE FORM

- Annual turnover of the bidder for the last 3 years ending March 2022

- HLL reserves right to accept/reject any/all bids at their discretion without assigning any reason.

- HLL reserves right to cancel this tender at any point of time without assigning any reason.

- The Qualified Vendors of HLL Lifecare Ltd shall stand qualified in the Technical Bid. Vendors who are new to HLL will undergo Vendor Development Protocol for the qualification of material, if they meet the minimum eligibility criteria. Vendors who do not meet the Minimum Eligibility Criteria will be disqualified and their Price bids will not be opened.

- QUANTITY:-

The Approximate Annual Requirement of Medical grade DEHP based PVC Sheet as per HLL Specification shall be 200 MT (2,00,000 Kg) per Annum. The quantity mentioned herein is approximate annual requirement and in case the company requires less/more quantity, the supplier should be prepared to affect supply at short notice at the agreed rate, terms and conditions.

- BIDDING PROCESS:-Interested prospective bidders can participate in the Invitation for Bid (IFB). Interested Bidder may also attend the Pre-bid Meeting as scheduled in the IFB, either On-Line or Off-Line to clarify the doubts and queries.  Bidder shall attach all  documents as required as per the Pre-Qualification Criteria in all respect  in the GeM Portal on or before of the last date of submission bid mentioned in the IFB and subsequent amendments/addendum if any. The Price Bids of bidders who fulfill the Prequalification Criteria only will be opened for further evaluation. HLL Lifecare Limited reserves the right to reject any or all of the offers without assigning any reason.

- PRICE BID (COMMERCIAL BID): The price quoted should be inclusive of Basic Price, Packing & Forwarding, Transportation, Insurance, other levies, etc and GST will be extra as applicable. HLL will not bear the cost of setting up of extrusion facility. The cost for setting up of the extrusion facility as per the requirements of HLL shall be amortized through the supply of sheet.

L1 Rate Bidder meeting all requirement of IFB will be successful Bidder. The quoted rate shall be valid for a period of One year, during that period no price escalation and no change in terms and conditions will be allowed on any ground. However when the raw material (PVC/DOP) price fluctuation (upward and downward changes) is more than 10% of market price, in such cases the rate will be amended as per the below indicated formula. The details of such changes will be included at the time of supply agreement.

The reference rate of PVC Resin/ DOP for this tender will be taken as per rate in approved polymer price portals such as www.polymerupdate.com, www.IndianPetrochem.com, etc. Bidder shall mention the material price as on the date of submission of bid. Rate will be reviewed on quarterly basis and whenever there is a rate increase/decrease by more than 10% of the rate already agreed for that material, the rate for supply of PVC Sheet, already agreed shall be refixed as per the formula given below:

PVC Resin

Current Rate of PVC Resin		= A (Rs)

Increased/Decreased Rate		= B (Rs)

Difference 				= B-A

% OF Increase /Decrease in rate 	= B-A X 100 = R

A

40% of R 				= S

Increase/ Decrease in rate 		= Existing Rate x (1+S%)

DOP

Current Rate of DOP 			= C (Rs)

Increased/Decreased Rate 		= D (Rs)

Difference 				= D-C

% OF Increase /Decrease in rate 	= D-C X 100 = X

C

20% of X 				= Y

Increase/ Decrease in rate 	= Existing Rate x (1+Y%)

Rates shall be refixed only when there is a +/- 10% variation from the rates already fixed/ refixed for that material.

- The Rates shall be filled in the BOQ of GeM Portal. The Price Bid shall be valid for a period of Five years from the date of opening of price bid. No price escalation and no change in terms and conditions will be allowed on any ground till the contract completed.

- The supplier shall specifically mention all Taxes/Duties/Levies in the Price Bid BOQ.

- ACCEPTANCE OF CONTRACT & PURCHASE ORDERS:-The successful Bidder has to execute an agreement with HLL Lifecare Limited, Akkulam Factory, Sreekaryam PO, Thiruvananthapuram, Kerala, which includes the Terms and Conditions of the Contract. The Agreement shall be as per the format given by HLL. Party also has to sign the duplicate copy of the Letter of Intent (LOI) / Purchase Order (PO) and return back within 7 days of receipt of LOI / PO for the supply of PVC Sheet, after the Setting Up of Extrusion Facility and Qualification of PVC Sheet by HLL, which will be considered as the confirmation of acceptance of LOI / PO.

- SECURITY DEPOSIT (SD):- Successful Bidder shall submit Security Deposit in Demand Draft for an amount of 3% of the Order Value, to a maximum of Rs 2,00,000 within 15 days of acceptance of order OR as per the Security Deposit Clause of HLL Purchase Manual (will be recovered from their running bills for the required security amount). The EMD submitted by the successful bidder shall be converted to Security Deposit and the bidder shall be allowed to remit the balance amount. In the case of Bidders under the category of MSE’s who have not submitted EMD, the security deposit will have to be remitted. SD will be recovered from the running Bills, to a maximum of Rs 2,00,000 of the supplier. The Security Deposit shall be payable to the Purchaser as compensation for failure to complete its obligations under the contract. Security Deposit will be discharged within 30 days from the date of completion of supply of tendered quantity.

- TIME LINE TO SET UP EXTRUSION FACILITY & START SUPPLY

The Time Line to set up the extrusion facility and start the trails, qualification and supply to HLL shall be 6 Months from the date of Letter of Intent / Agreement. Successful bidder shall HLL submit a detailed schedule of activities, which confined to maximum period of 6 Months to start supply the PVC Sheet as per the specification of HLL. Representatives / the third party deployed by HLL will visit the facility to review the progress of setting up of the Extrusion Facility.

- HLL shall issue Purchase Orders on staggered basis with Successful Bidders as per the requirement of HLL. The item should be delivered as per the delivery schedule mentioned in the Letter of Intent / Purchase Order.

- Acceptance of the delayed supplies and excess quantity is solely at the discretion of HLL. In the case of excess supply the excess quantity shall be returned back through the Transporter on freight to pay basis as indicated in Clause No 19.

- Loading of the item from suppliers site will be in the scope of tenderer. Unloading at the site of HLL will be borne by HLL.

- The Supplier shall supply strictly as per the order terms in respect of quality and specification.  HLL reserves the right not to receive the goods beyond the delivery date given in the order.

- In case of rejection the material should be taken back and replaced at supplier’s risk and cost within 7 days of intimation from HLL. The procedure in this regard would be as follows:-

- Material shall be returned through the Transporter indicated in Schedule F or through our approved Transporter on freight to-pay basis.

- Alternatively, the rejected material shall be destroyed at our end.

- The cost incurred for disposal of the rejected material shall be recovered from supplier.

- The supplier is required to replace the rejected material failing which HLL reserves the right to procure the material from alternate source(s) at supplier’s risk and cost.

- HLL reserves the right to inspect the facilities for GMP by their officials or by an authorized independent third party agency for compliance.

- All dispatch documents like delivery note/challan, packing list and invoice should contain the following details

- Supply Order No.& date

    - Description of items as contained in the supply Order.

    - Quantity dispatched

    - Manufacturing date

    - Expiry date

    - Total Number of packages/ serially numbered

    - Test certificate & COA

    - HLL’s GST No.

- TERMINATION OF THE CONTRACT

The Purchaser, without prejudice to any other contractual rights and remedies available to it (the Purchaser), may, by written notice of default sent to the Bidder, terminate the contract in whole or in part, if the Bidder fails to perform any other contractual obligation(s) within the time period specified in the contract, or within any extension thereof granted by the purchaser pursuant to Terms & Conditions of this Tender. In the event of the Purchaser terminates the contract in whole or in part, pursuant to Terms & Conditions, the Purchaser may procure services similar to those cancelled, with such terms and conditions and in such manner as it deems fit and the Bidder shall be liable to the Purchaser for the extra expenditure, if any, incurred by the Purchaser for arranging such procurement.

HLL reserves right to cancel the contract, if the successful bidder fails to adhere the contractual obligations such as,

- If the successful bidder fail to Execute the Agreement with HLL

- If the successful bidder fail to submit the schedule of setting up of the Extrusion Facility

- If the successful bidder fail to adhere the time line and delay in setting up of the Extrusion Facility

- If the successful bidder fail to supply the PVC Sheet after setting up of the Extrusion Facility as per HLL Requirement.

If the contract is terminated due to the any of the above reasons, the Security Deposit submitted by the party will be forfeited and they will be blacklisted as per the prevailing Purchase Rules. The decision of HLL will be final on this matter.  Unless otherwise instructed by the Purchaser, the Bidder shall continue to perform the contract to the extent not terminated.

- LIQUIDATED DAMAGES FOR DELAY:-The parties have to abide by delivery schedule given in the supply order      strictly. Penalty for delay @ 0.5% value of the materials per week of delay subject to a maximum of 7.5% of the value of the supply order will be imposed, if equipment is accepted by the Company after the stipulated delivery period.

- FORCE MAJEURE:- For purposes of this Clause “Force Majeure” means an event beyond the  control  of  the  Tenderer  and  not  involving  the  Tenderer’s  fault  or negligence and not foreseeable. Such events may include, but are not limited to, acts of the HLL either in its sovereign or contractual capacity, as or revolutions, fires,  floods,  epidemics,  quarantine  restrictions and freight embargoes. If a Force Majeure situation arises, the Tenderer shall promptly notify in writing of such conditions and the cause thereof. Unless otherwise directed by the Purchaser in writing, the Tenderer shall  continue to perform its obligations under the Contract as far as is reasonably practical, and shall seek all reasonable alternative means for performance not prevented by the force majeure event.

Notwithstanding the provisions contained in Terms & Conditions, the Bidder shall not be liable for imposition of any such sanction so long the delay and/or failure of the Bidder in fulfilling its obligations under the contract is the result of an event of Force Majeure.

For purposes of this clause, Force Majeure means an event beyond the control of the Bidder and not involving the Bidder’s fault or negligence and which is not foreseeable and not brought about at the instance of, the party claiming to be affected by such event and which has caused the non–performance or delay in performance. Such events may include, but are not restricted to, acts of the Purchaser either in its sovereign or contractual capacity, wars or revolutions, hostility, acts of public enemy, civil commotion, sabotage, fires, floods, explosions, epidemics, quarantine restrictions, strikes excluding by its employees , lockouts excluding by its management, and freight embargoes.

If a Force Majeure situation arises, the Bidder shall promptly notify the Purchaser in writing of such conditions and the cause thereof within seven days of occurrence of such event. Unless otherwise directed by the Purchaser in writing, the Bidder shall continue to perform its obligations under the contract as far as reasonably practical, and shall seek all reasonable alternative means for performance not prevented by the Force Majeure event.

If the performance in whole or in part or any obligation under this contract is prevented or delayed by any reason of Force Majeure for a period exceeding sixty days, either party may at its option terminate the contract Termination of this agreement for whatever reason shall not affect the obligation/ liabilities of both the parties accrued hereunder in respect of matters at the time of the agreement.

In case due to a Force Majeure event the Purchaser is unable to fulfill its contractual commitment and responsibility, the Purchaser will notify the Bidder accordingly and subsequent actions shall be taken on similar lines described in above sub- paragraphs.

- TERMINATION FOR INSOLVENCY:-The Purchaser may at any time terminate the Contract by giving written  notice to the Tenderer, if the Tenderer becomes bankrupt or otherwise insolvent. In this event, termination will be without compensation to the  Tenderer, provided that such termination will not prejudice or affect any right of action or remedy which has accrued or will accrue thereafter to the Purchaser.

- PAYMENT TERMS :-Within 30 days of Supply and Acceptance of the Materials at HLL Stores along with necessary documents.

- The tender is liable to be suspended or cancelled at any time at the discretion of the company without assigning any reason.

- RESOLUTION OF DISPUTES:- If dispute or difference of any kind shall arise between the Purchaser and the Bidder in connection with or relating to the contract, the parties shall make every effort to resolve the same amicably by mutual consultations.

If the parties fail to resolve their dispute or difference by such mutual consultations within twenty one days of its occurrence the same shall be referred for arbitration to a Sole arbitrator to be appointed by the C&MD, of the Purchaser Arbitration shall be conducted as per the provisions of the Arbitration and Conciliation Act 1996 or any statutory modification or re-enactment thereof for the time be in force. The arbitrator shall give a reasoned award. The award of the arbitrator shall be final and binding on the parties to this contract. The venue of arbitration shall be in the court of Thiruvananthapuram, Kerala, India. The jurisdiction of any disputed, suits and proceedings arising out of this IFB shall be only in the court of Thiruvananthapuram. The language of the proceedings shall be English.

एचएलएललाइफकेयरलिमिटेड/HLL LIFECARE LIMITED

भारतसरकारकाउद्यम/A GOVT.OF INDIA ENTERPRISE

आककुलमफ़ैक्टरि /Akkulam Factory, श्रीकारियम.पी.ओ। /Sreekariyam. P.O,

तिरुवनंतपुरम-६९५०१७ / THIRUVANANTHAPURAM-695 017

TENDER No: HLL/AFT/MTLS/EXTRUSION FACILITY/2022-23 			     10-01-2023

SECTION V

## SPECIAL CONDITION TO THE BIDDERS

Instructions to bidder are broad guidelines to be followed while formulating the bid and its submission to the Purchaser. It also describes the methodology for opening and evaluation of bids and consequent award of contract.

## DEFINITIONS

In this Contract, the following words and expressions shall have the meanings as stated below:

  - ‘Expression of Interest (EOI)’ shall mean and include the present document, and such other complements and agenda, which may subsequently be issued in this connection.

  - ‘Bidder / Tenderer’ shall mean the person, firm or Corporation submitting a bid against this invitation for bid and shall also include his agents and representatives.

  - ‘Purchaser / Owner’ shall mean HLL Lifecare Limited (HLL) (Akkulam Factory, Sreekariyam (PO), Thiruvananthapuram -17) or its units thereof.

  - ‘Officer-In-Charge'‚ shall mean the Officer appointed by the purchaser/owner to supervise all activities of the project.

  - ‘Bidder / Contractor'‚ shall mean the successful bidder whose bid has been accepted by the purchaser/owner and to whom the order is placed by the purchaser/owner and shall include his heirs, legal representatives, successors etc.

  - ‘Acceptance Letter'‚ shall mean written consent by a letter of purchaser/owner to the bidder intimating him that his bid has been accepted.

  - ‘Contract'‚ shall mean the articles of Contract Agreement, the conditions of contract, schedule of quantities, specifications, drawings attached and duly signed by the purchaser/owner and the Bidder.

  - ‘Date of Contract’‚ shall mean the date on which the successful bidder has accepted the notification of award.

  - ‘Contract Period’‚ shall mean the period (including rainy season) specified in the IFB documents during which the contract shall be executed.

  - ‘Completion Certificate'‚ shall mean the certificate   issued by the purchaser/owner to the Bidder after successful completion of the project.

## SCOPE OF THE BID

HLL Lifecare Limited (HLL), a Government of India Enterprise, (Akkulam Factory, Sreekariyam (PO), Thiruvananthapuram -17) invites Invitation for Bids (IFB)  from the eligible, competent and experienced prospective bidders who are capable of executing the Specified work as per our IFB conditions.

## ELIGIBLE BIDDERS

A Bidder should have the eligibility as per the Pre-Qualification (Minimum Eligibility Criteria) to submit bids against this IFB.

A firm/bidder shall submit only one bid in the same bidding process. A Bidder (either as a firm or as an individual or as a partner of a firm) who submits or participates in more than one bid will cause all the proposals in which the Bidder has participated to be disqualified.

## COST OF BIDDING

The Bidder shall bear all costs associated with the preparation and submission of its bid, and “the Purchaser", will in no case be responsible or liable for these costs, regardless of the conduct or outcome of the bidding process.

IFB documents may be downloaded free of cost from the Government e-Market Place GeM portal.

## INFORMATION OF THE INVITATION FOR BID (IFB)

All prospective bidders are expected to see all information regarding submission of bid for the Work published in the GeM Portal and the website of HLL, www.lifecarehll.com during the period from the date of publication of IFB and up to the last date and time for submission of bid. Non observance of information published in the website shall not be entertained as a reason for any claim or dispute regarding a tender at any stage.

Techno-Commercial Bids of the bidders participating in the Invitation of Bid (IFB) shall be submitted online in the GeM portal. No manual submission of bids shall be entertained for the bidders under any circumstances.

The Government e-Market Place (GeM portal) shall not allow submission of bids online after the stipulated date & time. The bidder is advised to submit the bids well before the stipulated date & time to avoid any kind of network issues, traffic congestion, etc. In this regard, the department shall not be responsible for any kind of such issues faced by bidder.

## BIDDING DOCUMENTS

Content of Bidding Documents

The bidding documents shall consists of the following unless otherwise specified

      - Section I		Notice Inviting IFB

      - Section II		Pre-Qualification Criteria (Eligibility Criteria)

      - Section III	Technical Specification

      - Section IV	Terms & Conditions

      - Section V		Special Condition to the Bidders

      - Section VI	Bid Security Form

      - Section VII	Indemnity Clause

      - Section VIII 	Declaration

      - Section IX	Acceptance Letter

      - Section X		Pre-contract Integrity pact

      - Section X		Price Bid Format

The Bidder is required to login to the e-procurement portal and download the listed documents from the website as mentioned in IFB. Bidder shall save it in     their system and undertake the necessary preparatory work off-line and upload the completed bid at his convenience before the closing date and time of submission.

The bidder is expected to examine carefully all instructions, Conditions of Contract, Contract Data, Forms, Terms, Technical Specifications, Bill of Quantities, Annexure and Drawings in the Bid Document. Failure to comply with the requirements of Bid Document shall be at the Bidder’s own risk.

- CLARIFICATION OF BIDDING DOCUMENT

A prospective bidder requiring any clarification of the bidding documents shall post through GeM Portal for clarifying the queries Tender Inviting Authority before the date of closing the Bid Date.

In case the clarification sought necessitates modification of the bid documents, being unavoidable, the Tender Inviting Authority may affect the required modification and publish them in the GeM through corrigendum.

## PREPERATION OF BIDS

Language of the documents relating to the IFB shall be in the English language.

Bidders shall not make any addition, deletion or correction in any of the bid documents. If tampering of documents is noticed during bid evaluation, the bid will be rejected and the bidder will be blacklisted.

- BID PRICES

Bidder shall quote their competitive rate in the format attached in GeM Portal. The Price Bids of tenderers who fulfill the Technical bid requirements only will be opened for further evaluation. The quoted rate shall be valid for a period of five years, during that period no price escalation and no change in terms and conditions will be allowed on any ground. However when the raw material (PVC/DOP) price fluctuation (upward and downward changes) is more than 10% of market price, in such cases the rate will be amended as per the Clause 8 of Terms & Conditions. The financial bid will have deliverable as per the requirement.

In the event of failure on the part of the supplier to supply the material as per the delivery schedule or supply material at the agreed rate during the contract period, HLL would procure the material from alternate source at the risk and cost of the supplier. In such cases, the additional cost incurred will be recovered from the supplier.

Failure to supply the material for any of the above reasons will also entail forfeiture of Security Deposit and even blacklisting of the supplier for a minimum period of two years.

- BID SECURITY (EMD)

The Bidder shall furnish, as part of his Bid, a Bid Security (EMD) as detailed in the IFB, Section IV, Clause 1 & Section I Clause 2, 3, 4 & 5. For GeM e-tenders, Bidders shall upload the payment details of Bid Security in the e-tender portal.

Each bid must be accompanied by Bid Security. Any Bid not accompanied by an acceptable Bid Security shall be rejected as non-responsive.

MSE units interested in availing exemption from payment of EMD should submit a valid copy of their Udyog Aadhaar registration certificate as mentioned in the NIT. But the Party has to provide Security Deposit if work is awarded to them. Security deposit will be 3 % of the order

- IFB DOCUMENT FEE

There is No Document Fee for this IFB

ALTERATIONS & ADDITIONS

The bid shall contain no alterations or additions, except those to comply with instructions, or as necessary to correct errors made by the bidder, in which case such corrections shall be initialed by the person or persons signing the bid.

The bidder shall not attach any conditions of his own to the Bid. The Bid price must be based on the IFB Terms & Conditions and Special Conditions to the Bidders. Any bidder who fails to comply with this clause will be disqualified.

- BID DOCUMENT SUBMISSION

The Bidder shall submit their bid online only through the Government e-Market Place (GeM portal) as per the procedure laid down for e- submission as detailed in the web site. For e tenders, the bidders shall download the IFB document from the portal. The Bidder shall fill up the documents and submit the same online in GeM Portal.

## QUALIFICATION CRITERIA

The bidder should be fulfilling the Prequalification Criteria as mentioned in Section II and preconditions and also must also upload/submit documentary evidence in support of fulfillment of these conditions while  submitting the bid.

HLL Lifecare Limited reserves the right to verify the credential submitted by the agency at any stage (before or after the award of contract). If at any stage, any information / documents submitted by the applicant is found to be incorrect / false or have some discrepancy which disqualifies the firm.

## DEAD LINE FOR SUBMISSION OF BIDS

Bid shall be received only online on or before the date and time as notified in IFB. The Tender Inviting Authority, in exceptional circumstances and at its own discretion, may extend the last date for submission of bids, in which case all rights and obligations previously subject to the original date will then be subject to the new date of submission. The Bidder will not be able to submit his bid after expiry of the date and time of submission of bid (server time).

## BID OPENING & EVALUATION

Bids shall be opened on the specified date, by the tender inviting authority or his authorized representative in GeM Portal

Date of Invitation of Bid (IFB) & Date of Opening of Technical Bid shall be mentioned in IFB. The intimation regarding acceptance / rejection of their bids will be intimated to the Bidders/firms through GeM portal. (shall contain scanned copy of Pre- qualification document.)

If any clarification is needed from bidder about the deficiency in their uploaded documents, they will be asked to provide it through GeM portal. The bidder shall upload the requisite clarification/documents within time specified by HLL Lifecare Limited, failing which bids will be liable for rejection.

In the event of the specified date of bid opening being declared a holiday for HLL, the bids will be opened at the same time on the next working day.

- CONFIDENTIALITY

Information relating to the examination, clarification, evaluation, and comparison of Bids and recommendations for the award of a contract shall not be disclosed to Bidders or any other persons not officially concerned with such process until the award has been announced in favour of the successful bidder.

Any effort by a Bidder to influence the Purchaser during processing of bids, evaluation, bid comparison or award decisions shall be treated as Corrupt & Fraudulent Practices and may result in the rejection of the Bidder’s bid.

- CLARIFICATIONS OF BIDS

To assist in the examination, evaluation, and comparison of bids, the Tender Inviting Authority may ask the bidder for required clarification on the information submitted with the bid. The request for clarification and the response shall be through GeM Portal.

No Bidder shall contact the Tender Inviting Authority on any matter relating to the submitted bid from the time of the bid opening to the time the contract is awarded. If the Bidder wishes to bring additional information to the notice of the Tender Inviting Authority, these documents shall be submitted along with the other documents.

During the bid opening, the Tender Inviting Authority will determine for each Bid whether it meets the required eligibility as specified in the IFB; is accompanied by the required Bid Security, documents and Certificates as required by the Prequalification Criteria.

A substantially responsive bid is one which conforms to all the terms, conditions, and requirements of the bidding documents, without material deviation.

Non submission of legible or required documents or evidences may render the bid non-responsive. Bidder can witness the principal activities and view the documents/summary reports for that particular work by logging on to the portal with his DSC from anywhere.

## AWARD OF CONTRACT

HLL will award the Contract to the Bidder whose bid has been determined to be substantially responsive and who has offered the lowest evaluated bid price for  Scope of Work mentioned in the Expression of Interest.

The Tender Inviting Authority reserves the right to accept or reject any Bid and to cancel the Bidding process and reject all Bids at any time prior to the award of Contract, without thereby incurring any liability to the affected Bidder or Bidders or any obligation to inform the affected Bidder or Bidders of the grounds for the Tender Inviting Authority‟s action.

## SECURITY DEPOSIT/PERFORMANCE SECURITY

Within thirty (30) days from date of the issue of notification of award by the Purchaser, the Bidder, shall furnish Performance Security to the Purchaser for an amount equal to three percent (3%) of the total value of the contract, valid up to 60 days after the date of completion of contract period / all contractual obligations by the Service provide.

The Performance security shall be denominated in Indian Rupees.

a)    It shall be in any one of the forms namely RTGS / NIFT. The validity of Bank Guarantee will be for a period up to sixty (60) days beyond respective contract Period.

In the event of any failure /default of the Bidder with or without any quantifiable loss to the Purchaser the amount of the performance security is liable to be forfeited. The Purchaser may do the needful to cover any failure/default of the Bidder with or without any quantifiable loss to the Purchaser.

In the event of any amendment issued to the contract, the Bidder shall, within twenty- one (21) days of issue of the amendment, furnish the corresponding amendment to the Performance Security (as necessary), rendering the same valid in all respects in terms of the contract, as amended.

The Bidder shall enter into contract agreement with the purchaser within 14 days upon issuance of notification of Award. The Purchaser will release the Performance Security without any interest to the Bidder on completion of the contract period including submission of satisfactory certificates received from Hospital authorities towards preventive maintenance and breakdown maintenance services rendered by the contractor during contract period.

## SUB CONTRACTS

The Bidder shall notify the Purchaser in writing of all sub contracts awarded under the contract if not already specified in its bid. Such notification, in its original tender or later, shall not relieve the Bidder from any of its liability or obligation under the terms and conditions of the contract. Sub contract shall be only for bought out items and sub-assemblies.

## PRICES

Prices to be charged by the bidder for performing the scope of work mentioned in the IFB in terms of the contract shall not vary from the corresponding prices quoted by the bidder during the contract period.

The variation in price on account of change in Rawmaterial cost shall be done as per Clause 8 of Terms & Conditions.

## TAXES & DUTIES

Bidder shall charge the taxes, duties, fees, levies etc. separately and shall be mentioned separately in the Price Bid. The Rate of these shall be as per the prevailing Rate.

## TERMS & MODE OF PAYMENT

Payment shall be made by RTGS within 30 days of receipt, inspection and acceptance of the material.

## DELAY IN BIDDERS PERFORMANCE

The Bidder shall perform the contract under the contract within the Timeline specified in the IFB and schedule given by the Purchaser and as incorporated in the contract.

Subject to the provision under Terms & Conditions, any unexcused delay by the Bidder in maintaining its contractual obligations towards performance of services shall render the Bidder liable to any or all of the following sanctions:

- Imposition of liquidated damages or

- Forfeiture of its performance security and

- Termination of the contract for default.

If at any time during the currency of the contract, the Bidder encounters conditions hindering timely performance of services, the Bidder shall promptly inform the Purchaser in writing about the same and its likely duration and make a request to the Purchaser for extension of the contract schedule accordingly. On receiving the Bidder’s communication, the Purchaser shall examine the situation as soon as possible and, at its discretion, may agree to extend the contract schedule, with or without liquidated damages for completion of Bidder’s contractual obligations by issuing an amendment to the contract.

When the period of contract is extended due to unexcused delay by the Bidder, the amendment letter extending the contract period shall, interalia contain the following conditions:

- The Purchaser shall recover from the Bidder, under the provisions of the clause of the Terms & Conditions, Liquidated Damages on the services, which the Bidder has failed to perform services stipulated in the contract.

- That no increase in price on account of any ground, whatsoever, including any stipulation in the contract for increase in price on any other ground and, also including statutory increase in or fresh imposition of applicable taxes or on account of any other tax or duty which may be levied in respect of services specified in the contract, which takes place after the date of contract period stipulated in the contract shall be admissible on such of the said services as are performed after the date of the contract period stipulated in the contract.

- But nevertheless, the Purchaser shall be entitled to the benefit of any decrease in price on account of reduction in or remission of applicable taxes or any other duty or tax or levy or on account of any other grounds, which takes place after the expiry of the service period stipulated in the contract.

## TERMINATION FOR DEFAULT

The Purchaser, without prejudice to any other contractual rights and remedies  available to it (the Purchaser), may, by written notice of default sent to the Bidder, terminate the contract in whole or in part, if the Bidder fails to perform any other contractual obligation(s) within the time period specified in the contract, or within any extension thereof granted by the purchaser pursuant to Terms & Conditions, Clause 21. In the event of the Purchaser terminates the contract in whole or in part, pursuant to Terms & Conditions, the Purchaser may procure services similar to those cancelled, with such terms and conditions and in such manner as it deems fit and the Bidder shall be liable to the Purchaser for the extra expenditure, if any, incurred by the Purchaser for arranging such procurement.

Unless otherwise instructed by the Purchaser, the Bidder shall continue to perform the contract to the extent not terminated.

## Termination for convenience

The Purchaser reserves the right to terminate the contract, in whole or in part for its (Purchaser’s) convenience, by serving 7days written notice on the Bidder at any time during the currency of the contract. The notice shall specify that the termination is for the convenience of the Purchaser. The notice shall also indicate interalia, the extent to which the Bidder’s performance under the contract is terminated, and the date with effect from which such termination will become effective.

The services which are complete and ready in terms of the contract for delivery and performance within thirty days after the Bidder’s receipt of the notice of termination shall be accepted by the Purchaser following the contract terms, conditions and prices. For the remaining services, the Purchaser may decide:

- To get any portion of the balance completed and delivered at the contract terms, conditions and prices; and/or

- To cancel the remaining portion of the services and compensate the Bidder by paying an agreed amount for the cost incurred by the Bidder towards the remaining portion of the goods and services.

## NOTICES

Notice, if any, relating to the contract given by one party to the other, shall be sent in writing or by E-Mail and confirmed in writing. The procedure will also provide the sender of the notice, the proof of receipt of the notice by the receiver. The addresses of the parties for exchanging such notices will be the addresses as incorporated in the contract.

The effective date of a notice shall be either the date when delivered to the recipient or the effective date specifically mentioned in the notice, whichever is later.

Joint General Manager (Materials)

एचएलएललाइफकेयरलिमिटेड/HLL LIFECARE LIMITED

भारतसरकारकाउद्यम/A GOVT.OF INDIA ENTERPRISE

आककुलमफ़ैक्टरि /Akkulam Factory, श्रीकारियम.पी.ओ। /Sreekariyam. P.O,

तिरुवनंतपुरम-६९५०१७ / THIRUVANANTHAPURAM-695 017

TENDER No: HLL/AFT/MTLS/EXTRUSION FACILITY/2022-23 			     10-01-2023

## SECTION VI

Bid Security Declaration Form

Date:	Tender No.

To

(insert complete name and address of the purchaser)

I/We. The undersigned, declare that:

I/We understand that, according to your conditions, bids must be supported by a Bid Securing Declaration.

I/We accept that I/We may be disqualified from bidding for any contract with you for a period of THREE years from the date of notification if I am /We are in a breach of any obligation under the bid conditions, because I/We

- have withdrawn/modified/amended, impairs or derogates from the tender, my/our Bid during the period of bid validity specified in the form of Bid; or

- having been notified of the acceptance of our Bid by the purchaser during the period  of bid validity (i) fail or reuse to execute the contract, if required, or (ii) fail or refuse to furnish the Performance Security, in accordance with the Instructions to Bidders.

I/We understand this Bid Securing Declaration shall cease to be valid if I am/we are not the successful Bidder, upon the earlier of (i) the receipt of your notification of the name of the successful Bidder; or (ii) thirty days after the expiration of the validity of my/our Bid.

Signature

(insert signature of person whose name and capacity are shown)

in the capacity of

(insert legal capacity of person signing the Bid Securing Declaration)

Name

(insert complete name of person signing he Bid Securing Declaration)

Duly authorized to sign the bid for an on behalf of (insert complete name of Bidder)Dated on	day of	(insert date of signing) Corporate Seal (where appropriate)

एचएलएललाइफकेयरलिमिटेड/HLL LIFECARE LIMITED

भारतसरकारकाउद्यम/A GOVT.OF INDIA ENTERPRISE

आककुलमफ़ैक्टरि /Akkulam Factory, श्रीकारियम.पी.ओ। /Sreekariyam. P.O,

तिरुवनंतपुरम-६९५०१७ / THIRUVANANTHAPURAM-695 017

TENDER No: HLL/AFT/MTLS/EXTRUSION FACILITY/2022-23 			     10-01-2023

SECTION VII

INDEMINITY CLAUSE

If the supplier fails to execute the order within the time prescribed for the delivery of goods ordered or violates or infringes the existing rates as agreed to as mentioned in the supply order, the supplier shall and will indemnify the company against all losses or damages whatsoever to be incurred or sustained including the legal cost or expenses incurred by the company by reason of non-delivery of goods at agreed quantity and rate with in the time specified in the supply order.  The company will initiate legal action if the supplier fails to execute the supply order as per the schedule in the supply order for the actual loss suffered. No quantity tolerance will be permitted in this regard unless otherwise prior approval is taken by the company before dispatching any excess quantity supplied which shall be returned back on freight to pay basis at the risk of the supplier. Responsiveness of the Bid shall be at the discretion of HLL.

The supplier shall have no right to change the quantity stipulated in the supply order.

Bid pronounced Non Responsive by HLL shall be summarily rejected.

The decision of HLL will be final and no correspondence on this shall be entertained.

We have read and understood the above conditions and agree to abide by the same.

PLACE:

DATE:                                                        	 NAME AND SIGNATURE OF THE APPLICANT

(WITH OFFICE SEAL)

एचएलएललाइफकेयरलिमिटेड/HLL LIFECARE LIMITED

भारतसरकारकाउद्यम/A GOVT.OF INDIA ENTERPRISE

आककुलमफ़ैक्टरि /Akkulam Factory, श्रीकारियम.पी.ओ। /Sreekariyam. P.O,

तिरुवनंतपुरम-६९५०१७ / THIRUVANANTHAPURAM-695 017

TENDER No: HLL/AFT/MTLS/EXTRUSION FACILITY/2022-23 			     10-01-2023

SECTION VIII

DECLARATION

We confirm having read and understood all the specifications, instructions, forms, terms and conditions and other requirements of the above tender (both expressed and implied) in full and that we agree to abide by all without any deviation.

PLACE:								SIGNATURE

DATE:							NAME AND ADDRESS OF APPLICANT

(SEAL OF THE APPLICANT)

एचएलएललाइफकेयरलिमिटेड/HLL LIFECARE LIMITED

भारतसरकारकाउद्यम/A GOVT.OF INDIA ENTERPRISE

आककुलमफ़ैक्टरि /Akkulam Factory, श्रीकारियम.पी.ओ। /Sreekariyam. P.O,

तिरुवनंतपुरम-६९५०१७ / THIRUVANANTHAPURAM-695 017

TENDER No: HLL/AFT/MTLS/EXTRUSION FACILITY/2022-23 			     10-01-2023

SECTION IX

ACCEPTANCE LETTER

IFB No.                                                             	dated

We, the undersigned have examined the above mentioned IFB document, including amendment/corrigendum No.	,dated	(if any), the receipt of which is hereby confirmed. We now offer to deliver (Description of services) in conformity with your above referred document, and agree to submit the Price Bid, if Technically qualified, as part of this IFB.

If our bid is accepted, we undertake to perform the services as mentioned above, in accordance with the contract.

We further confirm that, if our bid is accepted, we shall provide you with a Performance Security of required amount as per Terms & Conditions, read with modification. We agree to keep our IFB valid for acceptance, or for subsequently extended period, if any, agreed by us. We also accordingly confirm to abide by this IFB up to the aforesaid period and this bid may be accepted any time before the expiry of the aforesaid period. We further confirm that, until a formal contract is executed, this IFB read with your written acceptance thereof within the aforesaid period shall constitute a binding contract between us.

We further understand that you are not bound to accept the lowest or any bids you may receive against your above-referred IFB.

We confirm that we do not stand deregistered/banned/blacklisted by any Govt. Authorities. We confirm that we fully agree to the Terms and Conditions specified in above mentioned   IFB document, including amendment/ corrigendum if any

(Signature with date)

(Name and designation)

(SEAL OF THE APPLICANT)

एचएलएललाइफकेयरलिमिटेड/HLL LIFECARE LIMITED

भारतसरकारकाउद्यम/A GOVT.OF INDIA ENTERPRISE

आककुलमफ़ैक्टरि /Akkulam Factory, श्रीकारियम.पी.ओ। /Sreekariyam. P.O,

तिरुवनंतपुरम-६९५०१७ / THIRUVANANTHAPURAM-695 017

TENDER No: HLL/AFT/MTLS/EXTRUSION FACILITY/2022-23 			     10-01-2023

SECTION X

PRE-CONTRACT INTEGRITY PACT

- (In the company letter with sign & seal)

All bidders are bound to comply with the integrity pact clauses. Bids submitted   without signing the integrity pact will be ab initio rejected without assigning any reason.

HLL Lifecare Limited. Division :

Tender No:

- INTEGRITY PACT

This Pre-Contract Integrity Pact (herein after called the Integrity Pact) is made on

day of the month of

- Between

HLL Lifecare Ltd. a Government of India Enterprise with registered office at HLL Bhavan, Poojappura, Thiruvananthapuram 695 012, Kerala, India. (Hereinafter called “HLL”, which expression shall mean and include, unless the context otherwise requires, his successors in office and assigns) of the First Party.

- And

M/s	with   office   at   ……………………………………represented  by   Shri

,	Designation......................	(hereinafter	called	the

“Bidder/Seller”/Contractor/Agent which expression shall mean and include, unless the context otherwise requires, his successors and permitted assigns) of the Second Party.

Both HLL and Bidder referred above are jointly referred to as the Parties.

- Preamble

HLL intends to award, under laid down organizational procedures, Purchase orders / contract/s against Tender /Work Order /Purchase Order No.	HLL

desires full compliance with all relevant laws and regulations, and the principles of economic use of resources, and of fairness and transparency in its relations with its

Bidder/s and Contractor/s.

Page 34 of 41

NOW, THEREFORE,

To avoid all forms of corruption by following a system that is fair, transparent and free from any influence/prejudiced dealings prior to, during and subsequent to the currency of the contract to be entered into with a view to:-

  - Enable HLL to obtain the desired materials/ stores/equipment/ work/ project done at a competitive price in conformity with the defined specifications by avoiding the high cost and the distortionary impact of corruption on public procurement; and

  - Enable the Bidder to abstain from bribing or indulging in any corrupt practice in order to secure the contract by providing assurance to them that their competitors will also abstain from bribing and other corrupt practices and HLL will commit to prevent corruption, in any form, by its officials by following transparent procedures. The parties hereto hereby agree to enter into this Integrity Pact and agree as follows:

- Clause.1. Commitments of HLL

1.1 HLL undertakes that HLL and/or its Associates (i.e. employees, agents, consultants, advisors, etc.) will not demand, take a promise for or accept, directly or through intermediaries, any bribe, consideration, gift, reward, favour or any material or immaterial benefit or any other advantage from the bidder, either for themselves or for any person, organization or third party related to the contract in exchange for an advantage in the bidding process, bid evaluation, contracting or implementation process related to the contract.

  - HLL will, during the tender process / pre-contract stage, treat all BIDDERs with equity and reason, and will provide to all BIDDERs the same information and will not provide any such information or additional information, which is confidential in any manner, to any particular BIDDER which could afford an advantage to that particular BIDDER in comparison to other BIDDERs in relation to tendering process or during the contract execution.

  - All the officials of HLL will report to IEM, any attempted or completed breaches of the above commitments as well as any substantial suspicion of such a breach.

  - HLL will exclude from the process all known prejudiced persons and persons who would be known to have a connection or nexus with the prospective bidder.

  - If the BIDDER reports to HLL with full and verifiable facts any misconduct on the part of HLL’s Associates (i.e. employees, agents, consultants, advisors, etc.) and the same is prima facie found to be correct by HLL, necessary disciplinary proceedings, or any other action as deemed fit, including criminal proceedings may be initiated by HLL. Further, such an Associate may be debarred from further dealings related to the contract process. In such a case, while an enquiry is being conducted by HLL the proceedings under the contract would not be stalled.

Page 35 of 41

- Clause 2. Commitments of BIDDERs/ CONTRACTORs

- The BIDDER commits itself to take all measures necessary to prevent corrupt practices, unfair means and illegal activities during any stage of its bid or during any pre-contract or post-contract stage in order to secure the contract or in furtherance to secure it and in particular commit itself to the following:-

  - The BIDDER will not offer, directly or indirectly (i.e. employees, agents, consultants, advisors, etc.) any bribe, gift, consideration, reward, favour, any material or immaterial benefit or other advantage, commission, fees, brokerage or inducement to any official of HLL, connected directly or indirectly with the bidding process, or to any person, organization or third party related to the contract in implementation of the contract.

  - The BIDDER further undertakes that it has not given, offered or promised to give, directly or indirectly any bribe, gift, consideration, reward, favour, any material or immaterial benefit or other advantage, commission, fees, brokerage or inducement to any official of HLL or otherwise in procuring the contract or forbearing to do or having done any act in relation to obtaining or execution of the contract or any other contract with the Government for showing or forbearing to show favour or disfavor to any person in relation to the contract or any other contract with the Government.

  - The BIDDER will not engage in collusion, price fixing, cartelization, etc. with other counterparty(s).

  - The counterparty will not pass to any third party any confidential information entrusted to it, unless duly authorized by HLL.

  - The counterparty will promote and observe ethical practices within its  Organization and its affiliates.

  - BIDDER shall disclose the name and address of agents and representatives and Indian BIDDERS shall disclose their foreign principals or associates.

  - The counterparty will not make any false or misleading allegations against HLL or its Associates.

  - BIDDERs shall disclose the payments to be made by them to agents/brokers or any other intermediary, in connection with this bid/contract.

  - The BIDDER further confirms and declares to HLL that the BIDDER is the original manufacture/integrator/authorized government sponsored export entity of the defense stores and has not engaged any individual or firm or company whether Indian or foreign to intercede, facilitate or in any way to recommend to HLL or any of its functionaries, whether officially or unofficially to award the contract to the BIDDER, nor has any amount been paid, promised or intended to be paid to any such individual, firm or company in respect of any such intercession, facilitation or recommendation.

Page 36 of 41

  - The BIDDER while presenting the bid or during pre-contract negotiations or before signing the contract, shall disclose any payments he has made, is committed to or intends to make to officials of HLL or their family members, agents, brokers or any other intermediaries in connection with the contract and the details of services agreed upon for such payments.

  - The BIDDER will not accept any advantage in exchange for any corrupt practice, unfair means and illegal activities.

  - The BIDDER commits to refrain from giving any complaint directly or through any other manner without supporting it with full and verifiable facts.

  - If the BIDDER or any employee of the BIDDER or any person acting on behalf of the BIDDER, either directly or indirectly, is a relative of any of the officers of HLL,

BIDDER’s firm, the same shall be disclosed by the BIDDER at the time of filing of tender. The term ‘relative’ for this purpose would be as defined in Section 6 of the Companies Act 1956.

  - The BIDDER shall not lend to or borrow any money from or enter into any monetary dealings or transactions, directly or indirectly, with any employee of HLL.

  - The BIDDER will not collude with other parties interested in the contract to impair the transparency, fairness and progress of the bidding process, bid evaluation, contracting and implementation of the contract, and will not enter into any undisclosed agreement or understanding with other Bidders, whether formal or informal. This applies in particular to prices, specifications, certifications, subsidiary contracts, submission or non-submission of bids or any other actions to restrict competitiveness or to introduce cartelization in the bidding process.

  - The BIDDER will not commit any offence under the relevant Indian Penal Code, 1860 or Prevention of Corruption Act, 1988; further the Bidder(s)/ Contractor(s) will not use improperly, for purposes of competition or personal gain, or pass on to others, any information or document provided by the HLL as part of the business relationship, regarding plans, technical proposals and business details, including information contained or transmitted electronically. The BIDDER also undertakes to exercise due and adequate care lest any such information is divulged.

  - The BIDDER will not instigate third persons to commit offences outlined above or be an accessory to such offences.

  - The Bidder(s)/Contractors(s) of foreign origin shall disclose the name and address of the Agents/representatives in India, if any. Similarly the Bidder(s)/Contractors(s) of Indian Nationality shall furnish the name and address of the foreign Principal(s), if any.

  - The Bidder(s) shall not approach the courts while representing the matters to IEM and the Bidder(s) will await their decision in the matter.

- Page 37 of 41

- Clause.3. Previous contravention and Disqualification from tender process and exclusion from future contracts

  - The BIDDER declares that no previous contravention occurred in the last three years immediately before signing of this Integrity Pact, with any other company in any country in respect of any corrupt practices envisaged hereunder or with any Public Sector Enterprise in India or any Government Department in India that could justify BIDDER’s exclusion from the tender process

  - The BIDDER agrees that if it makes incorrect statement on this subject, BIDDER can be disqualified from the tender process or the contract, if already awarded, can be terminated for such reason. If BIDDER before award or during execution has committed a contravention through a violation of Clause 2, above or in any other form such as to put his reliability or credibility in question,  HLL is entitled to disqualify the BIDDER from the tender process.

- Clause.4. Equal treatment of all Bidders / Contractors / Subcontractors

4.1 The Bidder(s)/ Contractor(s) undertake(s) to demand from his Subcontractors a commitment in conformity with this Integrity Pact.

Bidders and Contractors.

4.3 HLL will disqualify from the tender process all bidders who do not sign this Pact or violate its provisions.

- Clause.5. Consequences of Violation / Breach

  - Any breach of the aforesaid provision by the BIDDER or any one employed by it or acting on its behalf (whether with or without the knowledge of the BIDDER) shall entitle HLL to take all or any one of the following action, wherever required:-

- To immediately call off the pre-contract negotiations without assigning any reason or giving any compensation to the BIDDER. However, the proceedings with the other BIDDER(s) would continue.

- If BIDDER commits violation of Integrity Pact Policy during bidding process, he shall be liable to compensate HLL by way of liquidated damages amounting to a sum equivalent to 5% to the value of the offer or the amount equivalent to Earnest Money Deposit/Bid Security, whichever is higher.

- In case of violation of the Integrity Pact after award of the contract, HLL will be entitled to terminate the contract. HLL shall also be entitled to recover from the contractor liquidated damages equivalent to 10% of the contract value or the amount equivalent to security deposit/ performance guarantee, whichever is higher.

- To immediately cancel the contract, if already signed, without giving any compensation to the BIDDER.

- To recover all sums already paid by HLL, and in case of an Indian BIDDER with interest thereon at 2% higher than the prevailing Prime Lending Rate of State Bank of India, while in case of a BIDDER from a country other than India with interest thereon at 2% higher than the LIBOR. If any outstanding payment is due to the BIDDER from HLL in connection with any other contract for any other stores, such outstanding payment could also be utilized to recover the aforesaid amount.

Page 38 of 41

- To encash the advance bank guarantee and performance guarantee /warranty bond, if furnished by the BIDDER, in order to recover the payments already made by HLL, along with interest.

- To cancel all or any other contract with the BIDDER. The BIDDER shall be liable to pay compensation for any loss or damage to HLL resulting from such cancellation/recession and HLL shall be entitled to deduct the amount so payable from the money(s) due to the BIDDER.

- To debar the BIDDER from participating in future bidding processes of HLL for a minimum period of five (5) years, which may be further extended at the discretion of HLL or until Independent External Monitors is satisfied that the Counterparty will not commit any future violation.

- To recover all sums paid in violation of this Pact by BIDDER(s) to any middleman or agent or broker with a view to securing the contract.

contract signed by HLL with the BIDDER, the same shall not be opened.

xi. Forfeiture of performance guarantee in case of a decision by HLL to forfeit the same without assigning any reason for imposing sanction for violation of the pact.

  - HLL will be entitled to all or any of the actions mentioned in para 5.1(i) to (x) of this pact also on the commission by the BIDDER or any one employed by it or acting on its behalf (whether with or without the knowledge of the BIDDER), of an offence as defined in Chapter IX of the Indian Penal Code, 1860 or Prevention of Corruption Act, 1988 or any other statute enacted for prevention of corruption.

  - The decision of HLL to the effect that a breach of the provisions of this Pact has been committed by the BIDDER shall be final and conclusive on the BIDDER. However, the BIDDER can approach the Independent External Monitor(s) appointed for the purposes of this Pact.

- Clause.6. Fall Clause

The BIDDER undertakes that it has not supplied/is not supplying similar product/systems or subsystems OR providing similar services at a price / charge lower than that offered in the present bid in respect of any other Ministry/Department of the Government of India or PSU and if it is found any stage that similar product/systems or sub systems was supplied by the BIDDER to any to the Ministry/Department of the Government of India or a PSU at a lower price, then that very price, with due allowance for elapsed time will be applicable to the present case and the difference in the cost would be refunded by the BIDDER to HLL, if the contract has already been concluded.

- Clause .7. Independent External Monitor(s)

  - HLL has appointed Sh. Ashok Kumar Mangotra IAS (Retd.) as Independent External Monitor(s) (hereinafter referred to as IEM(s)) for this Pact in consultation with the Central Vigilance Commission. Contact details of IEM is as below:

Sh. Ashok Kumar Mangotra IAS (Retd.), Independent External Monitor (IEM) Office: HLL Lifecare Limited, HLL Bhavan, Poojappura, Thiruvananthapuram 695 012, Kerala Email: iemhll@lifecarehll.com

Page 39 of 41

  - The responsibility of the IEM(s) shall be to review independently and objectively, whether and to what extent the parties comply with the obligations under this Pact.

  - The IEM(s) shall not be subject to instructions by the representatives of the parties and perform their functions neutrally and independently.

  - Both the parties accept that the IEM(s) have the right to access all the documents relating to the project/ procurement, including minutes of meetings.

  - As soon as the IEM(s) notices, or has reason to believe, a violation of this pact, he will so inform the CEO/CMD.

  - The BIDDER(S) accepts that the IEM(s) have the right to access without restriction to all project documentation of HLL including that provided by the BIDDER. The BIDDER will also grant the IEM(s), upon his request and demonstration of a valid interest, unrestricted and unconditional access to his project

The IEM(s) shall be under contractual obligation to treat the information and documents of the BIDDER/ Subcontractor(s) with confidentiality.

  - HLL will provide to the IEM(s) sufficient information about all meetings among the parties related to the Project provided such meeting could have an impact on the contractual relation between the parties. The parties will offer to the IEM(s) option to participate in such meetings.

  - The IEM(s) will submit a written report to the CEO/CMD of HLL within 8 to 10 weeks from the date of reference or intimation to him by HLL/BIDDER and, should consent arise, submit proposals for correcting problematic situations.

- Clause.8. Subcontractor(s)

If HLL obtains knowledge of conduct of a Bidder, Contractor or Subcontractor, or of an employee or a representative or an associate of a Bidder, Contractor or Subcontractor which constitutes corruption, or if HLL has substantive suspicion in this regard, HLL will inform the same to the Chief Vigilance Officer.

- Clause.9. Facilitation of Investigation

In case of any allegation of violation of any provisions of this Pact or payment of commission, HLL or its agencies shall be entitled to examine all the documents, including the Books of Accounts of the BIDDER and the BIDDER shall provide necessary information and documents in English and shall extend all possible help for the purpose of such examination.

- Clause.10. Law and Place of Jurisdiction

Both the Parties agree that this Pact is subject to Indian Law. The place of performance and hence this Pact shall be subject to Chennai Jurisdiction.

- Clause.11. Other legal Actions

The actions stipulated in the Integrity Pact are without prejudice to any other legal action that may follow in accordance with the provisions of the extant law in force relating to any civil or criminal proceedings.

Page 40 of 41

- Clause.12. Validity and Duration of the Agreement

This Pact begins when both parties have legally signed it. It expires for the Contractor/Successful bidder 12 months after the last payment under the contract or the complete execution of the contract to the satisfaction of the both HLL and the BIDDER /Seller, including warranty period, whichever is later, and for all other Bidders/unsuccessful bidders 6 months after the contract has been awarded.

If any claim is made / lodged during this time, the same shall be binding and continue to be valid despite the lapse of this pact as specified above, unless it is discharged / determined by Chairman and Managing Director of HLL.

- Clause. 13. Other provisions

13.1 Changes and supplements as well as termination notices need to be made in writing. Both the Parties declare that no side agreements have been made to this Integrity Pact.

13.1 If the Contractor is a partnership or a consortium, this agreement must be signed by all partners or consortium members.

13.1 Should one or several provisions of this agreement turn out to be invalid, the remainder of this agreement remains valid. In this case, the parties will strive to come to an agreement to their original intentions

INWITNESS THEREOF the parties have signed and executed this pact at the place and date first above mentioned in the presents of following witnesses:

- HLL Lifecare Ltd.							Bidder

Witness								Witness

1………………..…						1…………………

2……………………						2………………..

## Page 41 of 41

एचएलएललाइफकेयरलिमिटेड/HLL LIFECARE LIMITED

भारतसरकारकाउद्यम/A GOVT.OF INDIA ENTERPRISE

आककुलमफ़ैक्टरि /Akkulam Factory, श्रीकारियम.पी.ओ। /Sreekariyam. P.O,

तिरुवनंतपुरम-६९५०१७ / THIRUVANANTHAPURAM-695 017

TENDER No: HLL/AFT/MTLS/EXTRUSION FACILITY/2022-23 			     05-01-2023

SECTION XI

## PRICE BID FORMAT

SETTING UP OF EXTRUSION FACILITY AND SUPPLY OF

MEDICAL GRADE PVC SHEET FOR BLOOD BAG PRODUCTION

| Sl No | Description | 1stYear | 2ndYear | 3rdYear | 4thYear | 5th Year |
| --- | --- | --- | --- | --- | --- | --- |
| 1. | QUANTITY / YEAR IN MT |  |  |  |  |  |
| 2. | BASIC RATE / KG OF MEDICAL GRADE PVC SHEETFOR BLOOD BAG |  |  |  |  |  |
| 3. | PACKING & FORWARDING CHARGE/ KG |  |  |  |  |  |
| 4. | FREIGHT / KG |  |  |  |  |  |
| 5. | SUB TOTAL (1+2+3) |  |  |  |  |  |
| 6. | GST ………………… % |  |  |  |  |  |
| 7. | TOTAL RATE / KG (4+5) |  |  |  |  |  |
| 8. | TOTAL VALUE / YEAR ( 1 X 7 ) |  |  |  |  |  |

RATE OF PVC RESIN:…………………… as on Tender Submission Date ………………………

RATE OF DOP:…………………… as on Tender Submission Date ………………………

- The Price Bid shall be valid for a period of One Year and Contract Rate will be valid for five year of contract period.

- Rate will be firm & fixed during the contract period.

- Average Rate per Kg of Sheet will be taken for evaluation purpose.

- Rate of PVC Resin and DOP as on date of submission of bid shall be mentioned in the Price Bid.

PLACE:								SIGNATURE

DATE:							NAME AND ADDRESS OF APPLICANT

(SEAL OF THE APPLICANT)