##### HLL Lifecare Limited

TENDER NO: HLL/HCS/BLR/AMS/2013/15/LTE                           Date: 13/09/2013

ADDENDUM #1

25/09/2013

Sub: Addendum to the above referred Limited tender

With reference to the above, we hereby incorporate the following amendment to the tender documents as detailed below:-

Amendment No. 1

The Bid submission closing date and the bid opening date has been amended as below.

Bid submission closing date  :  30.09.2013,  12.00 PM

Bid opening date                   :  30.09.2013,  2.00 PM

Amendment No.2

The technical specification of the X-Ray machine inthe SCHEDULE -1  is been amended as follows.

SHEDULE :1

Specification for 800 mA  X-RAY Machine

| HIGH FREEQUENCY X-RAY MACHINE with 30 KHz and above, and approved by AERB |  |
| --- | --- |
| - X-RAY GENERATOR | Should be 800mA, 125 KV high frequency x-ray generator for general radiography. |
| - FREQUENCY OF GENERATOR | Should be 30 kHz and above |
| - OUTPUT POWER | Should be 50 KW |
| - KV RANGE | Should be 40 to 150 KVP |
| MA RANGE | Should be 800mA. |
| MAs RANGE | 1-200mAs or more |
| CONTROL | Should be a very compact, soft touch control panel having following functions & indications. The panel can be supplied in floor or wall mount and has a spill proof design. Following features Should be available on the control panel. - Machine ON/OFF switch. - Digital display of KV and mAs. - KV and mAs increase and decrease switches. - Ready and x-ray on switch with indicators - Bucky selection switch. A dual action hand switch with retractable cord shall be provided for radiation protection of operator. |
| X-RAY TUBE | Should be One no. of rotating anode, dual focus TOSHIBA/BHELL/ VARIAN/OPTILIX x-ray tube of 27/75 or more KW rating. The x-ray tube is thermally protected and has focal spot of 0.6 & 1.2 mm2 or better. Anode heat storage Should be capacity: 300 KHU |
| H.V. CABLE | Should be supplied with One pair of 8 meter H.V. cable compatible with the x-ray tube. |
| COLLIMATOR | Should be A manual, light beam diaphragm with knobs for adjustment of exposure light field. |
| TUBE STAND | Floor to ceiling stand and with counter balanced tube head (rotatable ± 180 degree) 360 degree rotatable; mounted on floor ceiling rails for convenient movements. |
| X-RAY TABLE | Should be of Floating type, table is with 4 way movement of the table top i.e. x axis and y axis. The table consists of motorized or non-motorized reciprocating bucky with grid of size 17 ¼” x 18 7/8” having grid ratio should be of 8:1 – 85 lines/inch & a stainless steel cassette tray. The bucky travels the entire length of the table and can be locked at any desired position by an electromagnetic lock. The tabletop is made of low radiation absorption, waterproof material. Compression band (immobilizing device) is provided with the table. |
| VERTICAL BUCKY STAND | Vertical bucky Should stand with oscillating grid of ratio 8:1, 85 lines/ inch. The bucky moves up and down and is should equipped with a stainless steel cassette tray. The stand is floor-mounted type and can accommodate cassettes up to 17 ¼” x 18 7/8”. Chest stand is provided to take chest radiographs, which is easily mountable / dismountable on the vertical bucky stand. The bucky can be tilted in 6 steps of 15 degree angle each for various radiographs. |
| - POWER SUPPLY | 3 Phase, 440Volts AC 50Hz with line resist less than 0.4 Ohms. Line Regulation  10%. |
| - The Machine should supplied along with the accesories | - LEAD APRON - 3 Nos - THYRIOD SHIELD - 2 Nos - Gonad shield - 2 Nos - Mobile lead barrier / permanent lead barrier -1 No Lead apron racks - For keeping 3 sets of Apron |

Group Head (HCS)