**Amendment/Addendum to Tender**

Sub: Tender for supply of Medical Equipments for Audiometry Room in the State of Haryana.

Ref: IFB NO: HLL/AFT-CMO/HR-AUDIO/EQP/2017-18 DATED 14.02.2018

The following addendum in the SECTION V (LIST OF REQUIREMENTS) , TECHNICAL SPECIFICATIONS and PRICE BID of the tender document referred above are issued:-

**SECTION V**

**LIST OF REQUIREMENTS**

|  |  |  |
| --- | --- | --- |
| **Sl. No** | **Equipments** | **Qty in Nos** |
| 1 | ABR screener | 13 |
| 2 | Audiometer | 1 |
| 3 | Portable Tympanometry Instrument | 12 |
| 4 | BERA with ASSR with both insert phone and head phone | 1 |
| 5 | Otoscope | 10 |
| 6 | Tympanometry (Model I) | 12 |

**Tympanometry (Model I)**

Ideal for basic screening tympanometry. Perform ipsilateral testing and get a simple overview of compliance curves and pass/fail results for your patient’s acoustic reflex.

It should able to do Ipsilateral test with colour display.

**Tehcnical specification:**

Probe frequency and intensity : 226 Hz ± 1% and 85 dB SPL in 2 ml ± 3 dB

Pressure range : -300 daPa to +200 daPa

**Volume range :** 0.1 ml to 6.0 ml ± 5%

**Integrated memory :** 100 tests

**Ipsilateral stimulation : Pure tone signals**

0,5 kHz, 1 kHz, 2 kHz, 3 kHz, 4 kHz ± 2%

**Probe level range**

80 dB to 110 dB SPL for all frequencies (5 dB steps)

**Physical characteristics:**

**Display :** 7.2” high resolution color display 800x480 pixel

**Probe cable length :** 3 meters

**Dimensions :** 330 mm x 220 mm x 55 mm (12.9 in x 8.6 in x 2.2 in)

**Weight :** Approx. 1.3 kg (2.9 lbs.)

**User interface :** English

**Standards:**

**Safety :** IEC 60601-1:1988 + A1:1991 + A2:1995 (edition 2.0)

ANSI/AAMI ES60601-1: A1:2012, C1:2009/(R)2012 and A2:2010/(R)2012

CSA CAN/CSA-C22.2 NO. 60601-1:14

IEC 60601-1 Edition 3.1 (2012)

Class II, Type B applied parts, IPX0

**EMC :** IEC 60601-1-2:2007 Edition 3.0 / Class B

**Performance :** IEC 60645-5:2004, Type 2

**SECTION – X PRICE SCHEDULE (A)**

**A) PRICE SCHEDULE FOR SUPPLY, INSTALLATION, TESTING AND COMMISSIONING OF EQUIPMENTS TO HARYANA**

| 1 | 2 | 3 | 4 |  | 5 | | | | | | | | 6 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| SI  No | Brief Description of Goods | Brand / Model | Quantity (Nos.) | Make | Price per unit (Rs.) | | | | | | | | Total Price (at Consignee Site) basis  (Rs.)  4 x 5(f) |
| Basic Price | GST | | Freight& Insurance & Delivery at Destination | | Service charges, if any, specify | Other charges, if any, specify | Unit Price (at Consignee Site) basis |
| **(f) =a+b+c+d+e** |
| (a) | (b) | | € | | (d) | € |
|  | % | Amt | % | Amt |
| 1 | ABR screener |  | 13 |  |  |  |  |  |  |  |  |  |  |
| 2 | Audiometer |  | 1 |  |  |  |  |  |  |  |  |  |  |
| 3 | Portable Tympanometry Instrument |  | 12 |  |  |  |  |  |  |  |  |  |  |
| 4 | BERA with ASSR with both insert phone and head phone |  | 1 |  |  |  |  |  |  |  |  |  |  |
| 5 | Otoscope |  | 10 |  |  |  |  |  |  |  |  |  |  |
| 6 | Tympanometry (Model I) |  | 11 |  |  |  |  |  |  |  |  |  |  |

Total tender price in rupees\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

**In words:** \_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

**Note:**

1. If there is a discrepancy between the unit price and total price THE UNIT PRICE shall prevail.
2. The charges for CAMC after warranty shall be quoted separately as per Section – X – Price Schedule (B)

**Name\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_**

**Business Address\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_**

**Place: \_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_ Signature of Tenderer\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_**

**Date: \_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_**

###### SECTION – X PRICE SCHEDULE (B)

1. PRICE SCHEDULE FOR COMPREHENSIVE ANNUAL MAINTENANCE CONTRACT AFTER WARRANTY PERIOD

|  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | 2 | 3 | 4 | | | | | | | |
| SI | BRIEF DESCRIPTION OF GOODS | Qty.  (Nos.) | Comprehensive Maintenance Contract price for Each Unit year wise\*. | | | | | | | |
| 1st year | 2nd year | 3rd year | 4thyear | 5th year | 6th year | 7th year | 8th year |
| a | b | c | d | e | f | g | h |
| 1 | ABR screener | 13 |  |  |  |  |  |  |  |  |
| 2 | Audiometer | 1 |  |  |  |  |  |  |  |  |
| 3 | Portable Tympanometry Instrument | 12 |  |  |  |  |  |  |  |  |
| 4 | BERA with ASSR with both insert phone and head phone | 1 |  |  |  |  |  |  |  |  |
| 5 | Otoscope | 10 |  |  |  |  |  |  |  |  |
| 6 | Tympanometry (Model I) | 12 |  |  |  |  |  |  |  |  |
|  | GST:\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_ | |  |  |  |  |  |  |  |  |
|  | Total Rs: | |  |  |  |  |  |  |  |  |

\* After completion of 1 year Warranty period for respective equipment as mentioned in the warranty clause

**NOTE:-**

1. In case of discrepancy between unit price and total prices, THE UNIT PRICE shall prevail.

2. The cost of CAMC which includes preventive maintenance including testing & calibration as per technical/ service /operational manual, labour and spares, after satisfactory completion of Warranty period may be quoted for next 8 years on yearly basis for complete equipment and Turnkey.

3. The cost of CAMC may be quoted along with taxes applicable on the date of Tender Opening. The taxes to be paid extra, to be specifically stated. In the absence of any such stipulation the price will be taken inclusive of such taxes and no claim for the same will be entertained later.

4. Successful tenderer shall enter into an agreement.

5. The payment of CAMC will be made as per clause GCC clause 21.1.

6. The uptime warranty will be 95 % on 24 (hrs) X 7 (days) X 365 (days) basis or as stated in Technical Specification of the TE document.

7. All software updates should be provided free of cost during CAMC period.

8. The stipulations in Technical Specification will supersede above provisions.

9. The supplier shall keep sufficient stock of spares required during Comprehensive Annual Maintenance Contract period. In case the spares are required to be imported, it would be the responsibility of the supplier to import and get them custom cleared and pay all necessary duties.

**Name\_**\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

**Business Address\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_**

**Place: \_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_ Signature of Tenderer\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_**

**Date: \_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_ Seal of the Tenderer\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_**

**All other terms and conditions in the original tender remain unchanged.**

**Dt. 09-04.2018                                                           Joint General Manager (Materials)**