**Date: 29/01/2015**

**Amendment No: 1**

**Ref: Tender Enquiry No: HLL/BME/Calibration-05/14-15 Dated 15.01.2015**

Pre-Bid meeting against the referred tender was held on 23/01/2015 at HLL Lifecare Ltd., Thycaud, Trivandrum- 695 014

The following amendments are issued with respect to the above Tender Enquiry Document.

**Technical Specification**

**Item No 18**

**Temperature Bath**

***Existing Specification:***

Para 2.2: Temperature range: -40°C to 125**°**C

***Amended as:***

Para 2.2: Temperature range: -30°C to 125**°**C

***Existing Specification:***

Para 2.3: Accuracy: 0.2%

***Amended as:***

Para 2.3: Accuracy 0.25**°**C

***Following lines are added to the existing specification:***

Para 2.6: Working area diameter should not be less than 40mm dia.

Para 2.7: Should be supplied with R type Thermocouple.

**All other contents, terms & conditions of the tender enquiry remain unchanged.**