Amendment - II

Tender for Design, Manufacture, Supply, Installation, Validation and Commissioning of

VENTILATOR AUTOCLAVE

Size: 1850 mm DIA X 4200 mm LONG

IFB No: HLL/AFT/PUR/AUTOCLAVE/2013-14	                                                                 DATE: 01.11.2013

As per clause No: 6.1 of Instruction to bidders, a Pre-Bid conference was held on 11/11/2013. The detailed discussion and clarification held with bidders are published herewith:

| Sl No. | Clause No | Queries raised | Clarification |
| --- | --- | --- | --- |
| 1 | Clause 31 of Instruction to bidders: Completion period | Needs 4 months for the delivery of the machine and one month for commissioning/handing over. This commissioning/handing over period considered is from the date of placement of machine at location and subsequent call for installation. | The bidder was asked to meet the stipulated timeline |
| 2 | Clause 30 of Instruction to bidders: Payment terms | a) 30% of the order value as advance payment against submission of equivalent advance BG. b) 60% within 7 days of dispatch. c) 10% against installation & commissioning or 60 days from the date of dispatch and against submission of equivalent Performance BG valid for 1year from the date of commissioning or 15 months from the date of shipment whichever is earlier. | HLL expressed that there shall not be any changes in the payment conditions available in tender document |
| 3 | Section V Technical Specification: Capacity | Bidder requires bag dimensions & loading details of the bag in the carriage. | The bag dimensions and loading details on carriages shall be shared with bidder. Drawings of SS 304 carriages available with HLL shall be shared with bidder. |
| Sl No. | Clause No | Queries raised | Clarification |
| 4 | Section V Technical Specification: Approved material of construction | Material of construction of Heat Exchanger fins & tube is SS316 instead of SS316L since all other contact parts are SS316L | Agreed. Material of construction of Heat Exchanger fins & tube shall be SS 316 |
| 5 | Section V Technical Specification: Inclusions | Instead of Aluminum mesh trays, it is suggested SS304. This is for ease of fabrication & longer life. | HLL indicated preference for perforated aluminum trays. However, in lieu of concerns raised by bidder on such trays, provision for quotes for both perforated aluminum trays as well as SS 316 trays has been given. |
| 6 | Section V Technical Specification: List of approved makes of component for Ventilator Autoclave | Jumo make Pressure Transmitter also be included in approved makes | JUMO shall also be included in list of approved makes along with SIEMENS |

Other Points:

1. Bidder shall provide as optional:

- 21 CFR compliance software for HMI for storing data

- SCADA system software.

2. Temperature probes:

Bidder shall provide additional temperature probe on top portion of chamber. Other probes will be as per the following distribution – one at bottom and two each at sides.

3. Peak demand requirement for all utilities such as compressed air, pure steam, plant steam to be indicated in the technical bid.

4. The bidder shall provide Soft Copy of PLC Programme in the form of CD.

Joint General Manager (Materials)